Is the wear of the Dual Mobility Cup inferior or superior of the conventional cup over a standard in vitro test?

Polyethylene (PE) wear is clearly linked to total hip arthroplasty (THA) failure, leading to osteolysis and decreasing survivorship rates. Dual mobility cups (DMC) are widely used to prevent THA instability. However, some studies have pointed DMC PE wear risk as a “dual wear” risk. Hip wear simulation is usually used to understand factors influencing wear and to differentiate design, PE and materials performances. To date, few works have been published studying dual mobility insert wear.

Our objective was to evaluate wear of DMC compared with a fixed single articulating hip design and to measure wear under same conditions (loading cycle, temperature, sterilization, material and surface roughness) approved by international standards ISO 14242-1 and ISO14242-2.

The test bench includes one station for a control sample and one for dynamic test. Those are driven independently one from the other. Two electrical actuators applied the forces and two forces sensors putted on the fixing plate of the acetabular part gave the corresponding values. On the dynamic station, the angular movements are generated by an electric motor. Sleeves are installed on the bowls containing the testing liquid and on the supports of acetabular parts, to get a tight volume that excludes contaminant particles. Wear is measured by a gravimetric method. The simulator is stopped and implants have been removed from the simulators to achieve weighting and observations at 0.5, 1, 2, 3, 4 and 5 million cycles. At the end, the sample PE insert and the control one are removed from their cup in the aim to measure the mass loss.

Under same conditions gravimetric wear and linear penetration of the head are perfectly comparable between a conventional and a dual mobility cup.

In vitro, DMC wear is equal or less important than a standard single fixed cup and volumetric wear is lower than published data. Wear of the two joints of a DMC is not increased thanks to the recruitment phenomenon and the freedom induced by the concept.

Bibliography:

1) Gaudin G., Ferreira A., Gaillard R., Prudhon J.L., Caton J.H., Lustig S. (2016) Equivalent wear performance of dual mobility bearing compared with standard bearing in total hip arthroplasty: in vitro study.

Int Orthop. doi : 10.1007/s00264-016-3346-5

2) Malatray M., Roux J.P., Gunst S., Pibarot V., Wegrzyn J. (2016)

Highly crosslinked polyethylene: a safe alternative to conventional polyethylene for dual mobility cup mobile component. A biomechanical validation.

Int Orthop. doi : 10.1007/s00264-016-3334-9

3) Laura A.D., Hothi H., Battisti C., Cerquiglini A., Henckel J., Skinner J., Hart A. (2016) Wear of dual-mobility cups: a review article.

Int Orthop. doi : 10.1007/s00264-016-3326-9

4) Philippot R, Boyer B, Farizon F, (2013)

Intraprosthetic dislocation: a specific complication of the dual-mobility system.

Clin Orthop Relat Res, 391 471(3):965-70.

5) Prudhon J-L, Ferreira A, Verdier R, (2013)

Dual mobility cup: dislocation rate and survivorship at ten years of follow-up.

Int Orthop, 37(12):2345-50.

6) Paxton EW, Inacio MC, Namba RS, Love R, Kurtz SM, (2015)

Metal-on conventional polyethylene total hip arthroplasty bearing surfaces have a higher risk of revision than metal-on-highly crosslinked polyethylene: results from a US registry.

Clin Orthop Relat Res, 473(3):1011-21.

7) Pineau V, Lebel B, Gouzy S, Dutheil J-J, Vielpeau C, (2010)

Dual mobility hip arthroplasty wear measurement: experimental accuracy assessment using radiostereometric analysis (RSA).

Orthops & Trauma Surgery & Research, 96(6):609-15.

Biography:

André FERREIRA, trained in LYON by Pr DEJOUR, has an expertise in TKA and THA which lead him to become vice chairman of SICOT HIP Committee. His works focus specially on implants fixation, new technologies and Dual Mobility cups that he has used for more than 20 years. He was one of the organizer of Lyon Hip Arthroplasty 2017 which made possible to update on them. He acquired this expertise through his practical experience, his research and his teaching.

He is General Secretary of French Society of Hip and Knee, associate editor of SICOT-J and correspondent member of International Orthoapedics, of Orthopaedic & Traumatology Surgery Review (OTSR) and of Revista de la Asociación Argentina de Ortopedia y Traumatología.

E-mail : a.ferreira@cliniqueduparclyon.com

www.chirurgie-hanche-genou.fr