An in vivo study of 3 artificial cervical discs: Are unconstrained designs susceptible to migration?

Jamie L. Baisden MD, Medical College of Wisconsin, Milwaukee, Wisconsin USA

Hoon Choi MD, Medical College of Wisconsin, Milwaukee, Wisconsin USA

Narayan Yoganandan PhD, Medical College of Wisconsin, Milwaukee, Wisconsin USA

Abstract

Statement of the Problem:  Artificial cervical discs (ACDs) have been used, instead of the conventional ACDF in the military. ACDs have potential advantages of motion preservation and reduction of adjacent segment de-generation.  Disc implants are classified as constrained, semi-constrained, and unconst-rained device types.  There have been independent reports of anterior migration and, sometimes disintegration, of non-constrained artificial discs.  In vivo studies investigating the potential migration of the artificial disc are limited.  Purpose:  The objective of this study is to examine the migration aspects in cervical discs using an in vivo model capable of simulating the axial load on the neck, and this mimics the human.  Methods:  C3-C4 ACD was performed with commercial implants in twelve adult Alpine Caprines. There were three groups: Group A received constrained; Group B received semi-constrained; Group C received unconstrained ACD.  Intra- and post-operative radiographs were performed to confirm hardware position.  Goats were monitored for six months postoperatively.  All goats tolerated cervical arthroplasty well and had satisfactory placement of the ACD implant on intra-operative radiography. There was no migration of the implants in Groups A and B.  Complete anterior extrusion of the inferior plate out of the disc space and disintegration of the implant occurred within one week after surgery in three goats in Group C. The fourth goat had anterior extrusion at five weeks.  Conclusions and Significance: Biomechanical responses under flexion/extension loading depend on the type of the ACD. Unconstrained ACD implants rely on axial loading and limited range of motion of the disc segment to preserve integrity. Unconstrained ACD implants may not be the most efficacious in individuals with a hypermobile neck and/or occupations involving head supported mass/helmets resulting in additional axial loads and repetitive motions of the cervical spinal column regions.

Recent Publications

- Alvin MD, et al. (2014) Cervical arthroplasty. The Spine Journal 14 (9):2231-2245.

- Ozbek Z, et al. (2017) Implant Migration in Cervical Disk Arthroplasty. World Neurosurgery 97:390-397.

- Tsermoulas G, Bhattathiri PS (2013) Anterior migration of prosthesis following cervical arthroplasty. British journal of neurosurgery 27 (1):132-133.

- Wagner SC, et al. (2014) Implant migration after Bryan cervical disc arthroplasty. The SpineJjournal, 14 (10):2513-2514.

- Wagner SC, et al. (2016) Traumatic Migration of the Bryan Cervical Disc Arthroplasty. Global spine journal 6 (1): e15-20.

Acknowledgments

Supported part by US DOD W81XWH-16-01-0110 and ZVAMC.

Biography

Jamie Baisden MD, is a Professor of Neurosurgery at the Medical College of Wisconsin specializing in Complex Spine. She has completed Fellowships in Complex Spinal Surgery and Spinal Cord Medicine and has a Masters in Medical Management from USC.  Her research interests include spinal biomechanics, spinal trauma, and evidence based medicine guideline development for spinal disorders.

Email: jbaisden@mcw.edu