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HPLC quantitative determination of Enalapril Maleate in 20 mg tablet.
Chem.Eng. Fatos Osmani, PhD. Melissa Troshani
A simple, rapid, economical, and highly sensitive stability-indicating HPLC method was developed and validated  for determination of enalapril maleate and related substances namely diketopiperazine.
Method; Chromatographic separation was achieved using an Agilent 1100 Series HPLC system and a column Zorbax C8 (250x4.6mm)(5-µm) and a analytical guard column Zorbax C8 (12.5x4.6mm)(5-µm) was connected to protect the column. 
Mobile phase:
A: 30mM H3PO4 adjusted at pH=2.0 with TEA+MeCN at ratio (65:35 v/v)
Flow rate: 1mL/min
UV Detection at 215 nm
CCT 40ºC
Dilution 40 ng/µL
RSD<1.5 %; correlation coefficient r2=0.99999;Resolution≥2.0; Tf<1.5; 
The method is fast and reliable and is used for routine quantitative analysis as well as quality control of enalapril maleate in pharmaceutical formulations, it is in use at Profarma Control Laboratories since May 2007.
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