Evaluation of performance of Deki Reader of Rapid Diagnostic Test for malaria diagnosis in rural military health facilities in Tanzania

Akili Kalinga1,7Charles Mwanziva3, Christopher Mswanya3, Deus Ishengoma4, Lucky Temu2, Lucas Mahikwano2, Saidi Mgata2, George Amoo7, Lalaine Anova5, Eyako Wurrapa5, Nora Zwingerman6,  Santiago Fero6, Geeta Bhat6, Ian Fine6, Mark Hickman5,   Colin Ohrt5, Reginald Kavishe8

1National Institute for Medical Research, Tukuyu Centre P.O. Box 538, Tukuyu, Tanzania

2 Henry Jackson Foundation Medical Research International, Dar es Salaam, Tanzania

3 Tanzania Peoples Defense Force, P.O. Box 9203 Dar es Salaam, Tanzania

4National Institute for Medical Research, Tanga Centre P.O. Box 5004, Tanga, Tanzania

5Walter Reed Army Institute of Research, 503 Robert Grant Ave, MD 20910, USA

# 6Fio Corporation, 111 Queen Street East, Suite 500, Toronto ON, M5C 1S2, Canada

7Amethysist Technologies LLC, 1450 S Rolling Road. Suite 4.0001 Baltimore, MD 21227, USA

# 8Kilimanjaro Christian Medical University College, P.O. Box2236, Moshi, Tanzania

Correspondence Author (contacts)

Akili Kalinga, Email: kalingaaka@yahoo.com; Phone: +255755380180

Abstract

Introduction: Although Microscopy is a standard diagnostic tool for malaria, is used at minimal with unreliable results because of unavailability of laboratory facilities in poor resource countries. Malaria Rapid diagnostic Tests (mRDTs) are currently advocated and used as adjunct to microscopy. However, at very low parasitaemia (<100p/µl) the test line on mRDT is very weak to be seen and consequently affecting interpretation of test results and patient care. Fio Corporation in Canada, developed a ruggedized portable, universal Deki Reader of mRDTs (DR of mRDTs) to perform automatic analysis and interpretation of RDT. However, before deploying the device for medical care in Tanzania, we evaluated its performance against microscopy as a reference test and compared to human interpretation of mRDTs

Methods: The cross-sectional study employed 1,293 outpatients with fever who were recruited and tested for malaria using mRDT and microscopy techniques.  Finger prick blood was prepared on mRDTs according to manufacturer’s instructions and test performed as guided DR of mRDT. Thick and thin blood smears were also prepared as guided by standardized template, stained and read by specialized Microscopist. We compared the performance of DR of mRDT to human interpretation of mRDT against Microscopy as gold standard

Results: Positivity rates by mRDT were 59.9% (775/1293) and 60.1% (777/1293) as interpreted by Human and DR respectively; parastaemia prevalence by microscopy as reference test was 48.4% (626/1293). The sensitivity of mRDTs interpreted by DR was 94.1% and that of manual interpretation was 93.9%. The specificity of DR of mRDT was 71.8% and that of human was 72.0 %. Positive Predictive Value (PPV) of mRDT by DR and Human was 75.8% and 75.4% respectively. The Negative Predictive Value (NPV) of mRDT by DR was 92.8% and by human was 92.4%.  There was no significance difference in sensitivity, specificity, PPV, NPV and accuracy of mRDT interpreted by DR and that of human interpretation

Conclusion: The performance of DR in interpreting mRDTs was found to be similar to human manual interpretation. There is a need to conduct more evaluations of performance of the device in different epidemiological settings and by using other type of mRDT assays for malaria diagnosis before validating its use in Tanzania