TITLE

DOES INFRARED TEMPERATURE REFLECT CORE BODY TEMPERATURE? A COMPARATIVE STUDY ON INFRARED TEMPERATURE FROM DIFFERENT PARTS OF THE BODY WITH AXILLARY TEMPERATURE.

Authors: 1. Dr Varghese Abraham MD, Consultant Paediatrician & In charge SNCU*

2. Mrs Elizabeth Johnson MSc, Nursing Superintendent*

3. Ms Khristina Deep GNM *

* Christian Hospital Chhatarpur, Chhatarpur, Madhya Pradesh, India (A Unit of Emmanuel Hospital  Association, New Delhi)

Correspondence Address: Dr Varghese Abraham MD

Consultant Paediatrician & In charge SNCU

Christian Hospital Chhatarpur, Mahoba

Road Chhatarpur

Madhya Pradesh, India – 471001

Email ID: varghesenettaabraham@gmail.com

ABSTRACT

Objective: To  determine the limit of agreement between Non-Contact Infrared Temperature (NCIT) from forehead, chest and abdomen with Digital Axillary Temperature(DAT).

Design: Prospective study

Setting: Neonatal unit of Christian Hospital Chhatarpur, Madhya Pradesh, India (A mission hospital under Emmanual Hospital Association, New Delhi)

Methods: In this study Non-Contact Infrared Temperature (NCIT) from the forehead, chest and abdomen was compared with Digital Axillary Temperature (DAT) by Bland Altman Plot.

Results: A total of 211 sets of temperature were recorded from neonates admitted in the Division of Neonatology, Christian Hospital Chhatarpur, Madhya Pradesh, India. The DAT agrees well with NCIT chest [mean difference =0.1346, 95% limit of agreement: (0.08455, 0.1846)] as compared to NCIT forehead and abdomen.

Conclusion: As compared to the tradition of using NCIT on the forehead, the author suggests the use of NCIT on the chest as it has more agreement to the core body temperature.

Key words: Temperature, Thermometer, Axillary, Infrared, Bland Altman