Abstract Template

Abstract Title: Comparative Study of a Novel Apnoea Event Detection Monitor with Portable Home Study Units

S.H. Kinsella, P.C. Loughlin, K.T Sweeney

1 First Author Affiliation (PMD Solutions Ltd. Ireland)

2Second Author Affiliation (University of Sydney, Australia)

3Third Author Affiliation (PMD Solutions Ltd. Ireland)

Abstract (not more than 300 words)

One of the key bottlenecks in the effective treatment of OSA is the lack of available clinics for sleep studies for definitive diagnosis. Waiting times for a sleep study varied widely in the early 2000's with referral to a specialist taking anywhere from 2 - 24 months. These inadequacies in recognition and diagnosis of OSA have resulted in an unacceptably high proportion of the population remaining undiagnosed, possibly up to 80% of people with OSA. These long wait times and the increased health and accident risks associated with undiagnosed OSA, has led to a strong push towards simpler, home-based methods for diagnosing (or 'ruling in') of patients with OSA. This study was designed to evaluate the sensitivity and agreement of PMD Solutions’ novel apnoea event screener compared to a Limited Home Sleep Test (HST). RespiraSense measures the movement of the chest and abdomen separately and simultaneously and transmits breath by breath data along with body positional data, wirelessly to a tablet device. The data from the study has a very good spread of Apnoea Hypopnea Index (AHI) values, ranging from Normal (0.3) to Severe (140+). This allowed for the algorithm to be tested across the entire range. Sensitivity (79%) and Specificity (87%) from the test results show that the RespiraSense sleep screener has significant promise in its ability to accurately screen a patient for sleep apnoea. This study has provided promising results towards the use of a non-invasive, take-home device which enables the identification of patients with suspected sleep apnoea. It is believed that the PMD Solutions Sleep Screener in conjunction with a questionnaire, such as STOP-BANG, can assist in tackling long waiting lists for Moderate-Severe apnoea suffers through quick and cost-effective screening.

Keywords—Sleep Apnea, Home Screening, Technology, Wearable

Professional Biography (100-150 words)

Seán Kinsella is a Medical Device professional with responsibility for the strategic direction for innovation, product development and clinical research at PMD Solutions. He holds an M.Sc. in Applied Physics in the area of robotics. Technical experience in hardware design, embedded software development and testing, sustaining of life supporting medical devices and 'Connected Health' medical device systems.

Dr. Patrick Loughlin completed his PhD at the University of Adelaide, South Australia, Australia with University of Adelaide with his thesis entitled  Elucidation of a peribacteroid membrane-bound bHLH transcription factor required for legume nitrogen fixation and is a past honorary research fellow at the University of Sydney, Australia. Dr. Kevin Sweeney completed his PhD. at the National University of Ireland, Maynooth in the area of Motion Artifact Processing Techniques for Physiological Signals.

# Recent Publications

| [1] | C. P. Subbe and S. H. Kinsella, “Continuous Monitoring of Respiratory Rate in Emergency Admissions: Evaluation of the RespiraSense™ Sensor in Acute Care Compared to the Industry Standard and Gold Standard,” Sensors, vol. 18, no. 8, p. 2700, 2018. |
| --- | --- |
| [2] | P. C. Loughlin, F. Sebat and J. G. Kellett, “Respiratory Rate: The Forgotten Vital Sign - Make It Count!,” The Joint Commission Journal on Quality and Patient Safety, vol. 44, no. 44, pp. 494-499, 2018. |

[3] Baker J.M., Riester C.J., Skinner B.M., Newell A.W., Swingley W.D., Madigan M.T., Jung D.O., Asao M., Chen M., Loughlin P.C., Pan H., Lin Y., Li Y., Shaw J., Prado M., Sherman C., Kuo-Hsiang Tang J., Blankenship R.E., Zhao T., Touchman J.W., and Sattley W.M. (2017) Genome sequence of Rhodoferax antarcticus ANT.BRT, a psychrophilic purple nonsulfur bacterium from an Antarctic microbial mat. Microorganisms (Accepted)

[4] Loughlin P.C., Duxbury Z., Mukasa Mugerwa T.T., Smith P.M.C., Willows R.D. and Chen M. (2016) Spectral properties of bacteriophytochrome AM1_5894 in the chlorophyll d-containing cyanobacterium Acaryochloris marina. Scientific Reports. 6, 27547

[5] Li Y., Lin Y., Garvey C.J., Birch D., Corkery R.W., Loughlin P.C., Scheer H., Willows R.D., Chen M. (2015) Characterization of red-shifted phycobilisomes isolated from the chlorophyll f-containing cyanobacterium Halomicronema hongdechloris. Biochimica et Biophysica doi: 10.1016/j.bbabio.2015.10.009

Paste your Photograph here

Email Id: sean.kinsella@pmd-solutions.com

Any Comments: