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Abstract (not more than 300 words)

Fluorescence has been used to monitor water quality for many years, initially within the laboratory and more recently in the field. The technique provides benefits in both sensitivity and specificity. Targeted compounds include Algae and Dye Tracing chemicals in the visible region of the spectrum and aromatic/heterocyclic hydrocarbons, tryptophan and optical brighteners at UV wavelengths. However, field fluorometers can be challenged by interfering fluorescence from non-target compounds, which can directly impact the accuracy of the readings obtained, which has limited their use. Further, non-linearity in the response can arise when either concentrations of the targeted compounds reach a level where absorption become significant, or turbidity levels are high, which limits the range of standard in-situ fluorometers. 

To overcome these issues, CTG has developed the V-Lux multiparameter fluorometer. This in-situ sensor provides 3 fluorescence channels combined with optical absorbance and turbidity measurements. The latter two parameters are used to correct the fluorescence readings to provide a linear response over an extended dynamic range. The provision of three fluorescence channels allows non-specific background interferences to be assessed and eliminated. 

A new calibration methodology has also been developed that reports fluorescence output intensity relative to a traceable Quinine Sulphate fluorescence standard. This approach ensures that the outputs from all fluorometer channels can be compared directly, without reference to the specific calibration compounds used. The approach also allows different fluorometers to be compared directly, which has not been possible up until now and has inhibited the wider uptake of fluorescence as a monitoring tool.
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