Comparison study between Light-Emitting-Diode and Quartz-tungstin Halogen C Light curing systems

ABSTRACT
 Objective: The purpose of this study was to measure the temperature rise caused by 10 light-emitting diode (LED) and three quartz-tungsten-halogen (QTH) curing lights and to relate the measured temperature rise to the power density of the curing light. The hypothesis was that the temperature rise induced by the LED curing lights would be smaller than that generated by the QTH curing lights operating at the same power density.
Materials and Methods: The curing lights evaluated in this study included a conventional QTH light (XL3000, 3M ESPE, St. Paul, MN, USA), two high-intensity QTH lights (Optilux 501, SDS/Kerr, Orange, CA, USA; Elipar Highlight, 3M ESPE), and 10 LED curing lights (Aqua Blue, Toesco, Kanagawa, Japan; CoolBlu, DentalSystems, Tokyo, Japan; DioPower, CMS, Copenhagen, Denmark; Elipar Freelight, 3M ESPE; Elipar Freelight 2, 3M ESPE; L.E. Demetron 1, SDS/Kerr; Lux-OMax, Akeda, Lystrup, Denmark; Lux-O-Max P1, Akeda; SmartLite, Dentsply, Konstanz, Germany; Ultra-Lume 2, Ultradent Products, South Jordan, UT, USA). The estimated relative heat generated by the curing lights was determined by measuring the temperature rise on the surface of a 4-mm × 4-mm cylinder of resin composite (Tetric Ceram; Ivoclar Vivadent, Schaan, Liechtenstein) using a thermocouple connected to a galvanometer (Radiometer, Copenhagen, Denmark). All measurements were made at the end of a 20-second exposure to the curing light. Power density was measured using a dental radiometer (Demetron Research, Danbury, CT, USA). The data were evaluated statistically by regression analysis and by the Neuman–Keuls multiple comparison test.
Results: Taking all curing lights into consideration, the coefficient of correlation between temperature rise and power density was very good and statistically significant at r = 0.93 (p < 0.001). The correlation coefficient for only the LED curing lights was even stronger, r = 0.96 (p < 0.001). This means, of course, that temperature rise increases with increased power density. On the plotted LED regression line, the temperature rise of two of the three QTH curing lights exceeded (by 2 and 2.3°C) the projected temperature rise of an LED curing light measured at the same power density (650mW/cm2). The measured temperature rise of the conventional QTH curing light operating at 360mW/cm2 was not significantly different from the predicted LED temperature rise regression line.
[bookmark: _GoBack]Conclusion: In general, the hypothesis that temperature rise caused by LED curing lights would be less than that generated by QTH curing lights operating at the same power density could not be confirmed. The data did not support earlier findings that LED curing units generated smaller temperature rises than QTH curing lights. The authors concluded that the main reason for earlier findings was that first-generation LED curing lights had lower power densities than present-day lights.
