## Slide 1
Non-invasive Measurement of Tympanic Membrane Oscillations
Hansen Mansy, Leslie Simms, Sierra Condo, Antonio Miranda, Kim Manwaring, Richard Sandler
University of Central Florida – Department of Mechanical and Aerospace Engineering, and College of Medicine
Continuous monitoring of intracranial pressure (ICP) in patients with hydrocephalus is crucial for the effective maintenance of brain shunt systems.
Recent studies suggested that there is a correlation between ICP and the oscillation of the tympanic membrane. Measurement of tympanic membrane oscillations may provide a non-invasive method for monitoring ICP.
Membrane oscillations may be measured by a variety of vibration sensors such as microphones and optical transducers. The resulting signals can be continuously displayed and analyzed using digital signal processing methods to quantify the different signal characteristics.
Introduction
Results
Acknowledgements
Future Studies
Conclusions
Financial support of for project was provided by NIH/NIBIB R01 EB012142
Two sensors were utilized
Piezoelectric sensor (Fig 1 and 2)
Converts pressures in the ear canal to electrical charge.
Materials & Methods
Figure 2: Piezoelectric ear piece.
Figure 1: Piezoelectric sensor head set in use.
Optical Sensor Calibration
Tympanic membrane movement (Fig 3) detected with the piezoelectric sensor, synchronized with the jugular vein pulses (Fig 4) measured by an accelerometer.
Optical Sensor Calibration (Cont)
Piezoelectric sensor is capable of detecting movement of the tympanic membrane, which is synchronous with the heart pulse.
Optical sensor may be capable of measuring tympanic membrane displacement due to its sensitivity. It was calibrated against a laser vibrometer.
Compare the piezoelectric and optical sensors.
Explore Laser Doppler Vibrometry as a method of measuring tympanic membrane behavior.
Quantify the correlation between tympanic membrane oscillations and ICP.
Optical sensor
Composed of a light emitting diode (LED) and photoresistor (placed side-by-side).
The LED emits light onto a surface and the photoresistor detects the reflected light and converts the light intensity to an electrical signal.
Materials & Methods cont.
References
Manwaring KH, Manwaring J: “Non-invasive ICP and compliance monitoring to detect shunt malfunction” Winter Meeting; Pediatric Section of the American Association of Neurosurgeons Salt Lake City, Utah / December 3, 2003
Gwer, S., Sheward, V., Birch, A., Marchbanks, R., Idro, R., Newton, C. R., & ... Lim, M. (2013). The tympanic membrane displacement analyser for monitoring intracranial pressure in children. Child's Nervous System: Chns: Official Journal Of The International Society For Pediatric Neurosurgery, 29(6), 927-933. doi:10.1007/s00381-013-2036-5
Figure 3: Piezoelectric sensor detects change in voltage from within the ear canal.
Chart: Time (seconds)
Chart: Time (seconds)
Figure 4: Accelerometer detects heart beat via blood pulsations of the external carotid artery.
Figure 15: Future studies include quantifying the correlation between tympanic membrane oscillations and the increase in ICP (induced using a tilt table).
Chart: Time (s)
Chart: LED & Shaker Minimum Amplitude 1Hz
LED & Shaker
Maximum Amplitude
1Hz
Chart: LED & Shaker (Amplified) Minimum Amplitude 1Hz
Figure 9: Optical Sensor Voltage Output at Maximum Shaker Amplitude
Figure 10: Optical Sensor Voltage Output at Minimum Shaker Amplitude
Figure 11: Amplified and Soldered Optical Sensor Voltage Output at Minimum Shaker Amplitude
Optical Sensor: Elementary Version
Optical Sensor: Improved Version
Chart: Laser vibrometer Velocity
 | 0.12194 | 0.22283 | 0.33611 | 0.55628 | 0.76911 | 0.86522 | 0.90616
Chart: Laser vibrometer Displacement (Integrated velocity)
 | 0.018267 | 0.01887 | 0.019062 | 0.01901 | 0.018822 | 0.018113 | 0.016543
Chart: LED Freq-RMS
| 0.86430423509075194 | 1.329433661260303 | 2.2522522522522523 | 4.6162723600692441 | 6.8385060494476591 | 7.7230359520639151 | 8.2206311323256429
 | 0.00091128 | 0.00091128 | 0.00092908 | 0.00092511 | 0.00089234 | 0.00077079 | 0.00075559
Figures 7 & 8: Shaker used to calibrate sensor
Figure 5: Elementary vs Improved Light Emitting Diode Sensor
Figure 6: Diagram of Ear Structure
Figure 12: Laser Vibrometer Frequency vs Velocity (RMS)
Figure 13: Laser Vibrometer Frequency vs Displacement (RMS )
Figure 14: Optical Sensor Frequency vs its Voltage output (RMS)