Introduction

Electrocardiograph machines and clinicians can recognize many common lead misplacements. However,

errors in arrangement of bipolar leads potentially create the unique predicament of pseudoinfarct ECG

patterns. Discerning a limb lead misplacement from a true myocardial infarction is challenging and,

unless there are a high clinical suspicion and an astute reader, is frequently overlooked. Here we

illustrate a case of left arm (LA)/left leg (LL) lead reversal, and briefly review the mechanism and clinical

importance of this under-recognized entity.

Case Report

A 91 year old hypertensive male, in otherwise excellent physical health, was evaluated for an isolated

syncopal episode. His vital signs and physical examination were unremarkable. A routine

electrocardiogram (ECG) was initially read as sinus rhythm, LVH with repolarization abnormality and

inferior infarct (new as compared to old ECG from 2013, Figure 2). An echocardiogram later revealed

results within normal limits. Further analysis by expert electrocardiogram readers found that the

(pseudo)inferior wall infarct pattern was due to a LA/LL lead reversal.

Discussion

Although in our case this error was quickly identified and rectified (Figure 3) this type of LA/LL reversal

has the potential to create confusion and in the right clinical setting can lead to expensive, unnecessary

and risky investigations and treatment leading to patient anxiety due to misinterpretation as an infarct.

The standard Einthoven’s triangle is represented in Fig. 1. In our patient due to LA/LL reversal the lead III

would be upside down. This leads to a pseudo inferior wall infarct pattern on the surface 12-lead ECG.

Frequently this error will lead to a near isoelectric line lead I, II or III, which might be the only sign of

LA/LL lead reversal in routine clinical practice. Computerized algorithms often lag in the automatic

recognition of this type of lead misplacement underlying the need for recognition of this subtle error.