# Marginal discrepancy of five contemporary dental ceramics for anterior restorations

Noor Nawafleh, PhD, Jordan University of Science and Technology, Jordan

### Abstract (300 word limit)

Statement of the problem: Over the last decades, significant efforts have been made to increase strength and reliability of the allceramic systems. Materials with esthetic advantages such as lithium disilicate ceramics have been very popular for the replacement of anterior teeth. More recently translucent zirconia and zirconia reinforced lithium silicate materials have been introduced as stronger materials with comparable esthetic capabilities. Purpose: This study aimed to measure and compare marginal accuracy of allceramic crowns indicated for anterior restorations.

Materials and Methods: A master die of maxillary central incisor was prepared for allceramic crown restoration and duplicated to produce 50 replicas of epoxy resin die material. Five ceramic materials (e.max CAD, Vita Suprinity, Celtra Duo, Cercon xt and e.max ZirCAD) were used to mill the crowns with ten specimens each. All crowns were manufactured following the same digital workflow; same master die, scanning unit and design software, and the recommended manufacturing protocol. Final seating of crowns was secured by a small droplet of temporary cement on its incisal edge. Marginal accuracy was evaluated by Scanning Electronic Microscope with a magnification of 300x. Vertical marginal gap was measured for each crown at predefined 4 points. One-way and Tukey multiple comparison tests were used to analyze significant differences between test groups (P= 0.05).

Results: The highest mean marginal gap and average maximum gap calculated were for the e.max CAD crowns (49.2 µm, 87.6 µm) while the lowest values were for the Cercon xt crowns (10.2 µm, 21.7 µm). The mean marginal gap and the average maximum gap of the e.max CAD crowns were statistically significantly greater than those of all other groups (p<0.05). However, the differences between all other combinations were insignificant (p>0.05).

Conclusions: Zirconia and zirconia reinforced lithium silicate materials can produce a greater level of marginal accuracy compared to lithium disilicate.

### Recent Publications (minimum 5)

- Nawafleh, N., Bibars, A. R., Al Twal, E., & Öchsner, A. (2020). Influence of antagonist material on fatigue and fracture resistance of zirconia crowns. European journal of dentistry, 14(02), 200-205

- Nawafleh, N., Bibars, A. R., Elshiyab, S., & Janzeer, Y. (2020). In vitro Simulation of Periodontal Ligament in Fatigue Testing of Dental Crowns. European Journal of Dentistry, 14(03), 380-385.

- Elshiyab, S. H., Nawafleh, N., Khan, U., Öchsner, A., & George, R. (2020). Impact of Coping Veneering Techniques on the Survival of Implant-Supported Zirconia-Based-Crowns Cemented to Hybrid-Abutments: An-In-Vitro Study. Bioengineering, 7(4), 117.

- Nawafleh, N., Hatamleh, M. M., Öchsner, A., & Mack, F. (2018). The impact of core/veneer thickness ratio and cyclic loading on fracture resistance of lithium disilicate crown. Journal of Prosthodontics, 27(1), 75-82.

- Elshiyab, S. H., Nawafleh, N., Öchsner, A., & George, R. (2018). Fracture resistance of implant-supported monolithic crowns cemented to zirconia hybrid-abutments: zirconia-based crowns vs. lithium disilicate crowns. The journal of advanced prosthodontics, 10(1), 65-72.

- 5. Elshiyab, S. H., Nawafleh, N., Walsh, L., & George, R. (2018). Fracture resistance and survival of implant‐supported, zirconia‐based hybrid‐abutment crowns: Influence of aging and crown structure. Journal of investigative and clinical dentistry, 9(4), e12355.

Paste your Logo here

Biography (150 word limit)

Noor Nawafleh is an assistant professor in Dental Technology at Jordan University of Science and Technology, Jordan. She graduated from the same university in 2002 and awarded a scholarship for being the top student in that year. Noor finished her Master and PhD at Griffith University, QLD, AUS. She earned Karl Ring Bursary for her outstanding academic performance. She has her expertise in Dental Materials, more specifically, ceramics.

Noor is very passionate to teaching and research. Since her graduations and during her study for Master and PhD degrees, she was working as a teaching and research assistant to enrich her skills and knowledge. She uses her positive attitude and tireless energy to encourage others, especially students, to work hard and succeed. Noor authored and co-authored many publications in prestigious journals focused in investigating fatigue behavior and oral environment simulation in testing ceramic materials.

Email: nanawafleh@just.edu.jo

Notes/Comments:

Presenting author details

Full name: Noor Nawafleh

Contact number: +962772758037

Mail id:

Category: (Poster presentation)