# Home self-monitoring of intraocular pressure using the ICare Home Rebound Tonometer

Jeanne J Ogle, MD, FRCS Edin

Department of Ophthalmology, Tan Tock Seng Hospital, Singapore

### Abstract

Purpose: To establish whether patients can perform accurate and reliable home self-monitoring of their intraocular pressures (IOPs) using the ICare Home.

Methods: 20 patients had their IOPs measured by the study team using the Goldmann applanation tonometer (GAT) and the ICare Home. They were then trained to use the ICare Home which was loaned to them for a period of 7 days. They measured IOPs for both eyes twice a day, taking 3 readings each time. A week later, GAT & ICare readings were again taken by the study team and the patients. The IOP readings were downloaded and analysed. Bland-Altmann difference analysis was used to assess agreement between GAT and ICare Home. Intraclass correlation (ICC) compared ICare Home measurements by the study team and patients, and test-retest variability for the patient’s readings over the 7-day period were analysed by taking the standard deviation of the 6 readings taken per day.

Results: There was good agreement between GAT and ICare Home readings obtained by the study team for the left eye (mean IOP difference 1+3.70 mm Hg, P=0.242) but less so for the right eye (1.75+3.32mm Hg, P=0.029), with ICare Home readings slightly higher for both eyes. Intraclass correlation coefficient (ICC) between ICare Home readings obtained by the study team and the patients at the first visit was 0.759 (95% confidence interval (CI), 0.474-0.900) for the right eye, and 0.783 (95% CI, 0.529 – 0.908) for the left eye. ICC at the final visit was 0.784 (95% CI 0.532-0.909) for the right eye, and 0.833 (95% CI 0.617-0.932) for the left eye, showing good agreement. Test-retest variability did not show any significant trend over the 7-day period.

Conclusions: Glaucoma patients are able to perform reasonably accurate self-tonometry using the ICare Home tonometer although the readings are generally higher compared to GAT.

Recent Publications

- Fernandes P, Diaz-Rey JA, Queiros A et al. Comparison of the ICare rebound tonometer with the Goldmann tonometer in a normal population. Ophthal. Physiol. Opt. 2005; 25:436-440.

- Abraham LM, Epasinghe NCR, Selva D et al. Comparison of the ICare rebound tonometer with the Goldmann applanation tonometer by experienced and inexperienced tonometrists. Eye 2008; 22: 530-506.

- Asrani S, Chatterjee A, Wallace DK et al. Evaluation of the ICare Rebound Tonometer as a Home Intraocular Pressure Monitoring Device. J. Glaucoma 2011; 20(2):74-79.

- Dabasia PL, Lawrenson JG, Murdoch IE. Evaluation of a new rebound tonometer for self-measurement of intraocular pressure. Br J Ophthalmol 2016; 100(8):1139-43.

- Mudie LI, LaBarre S, Varadaraj V et al. The Icare HOME (TA022) Study. Performance of an Intraocular Pressure Measuring Device for Self-Tonometry by Glaucoma Patients. Ophthalmology 2016;123(8):1675-84

Biography

Dr Jeanne Ogle is a Principal Resident Physician at the Department of Ophthalmology, Tan Tock Seng Hospital. She obtained her Fellowship from the Royal College of Surgeons (Edinburgh) in 1999 and Masters degree in Surgery (Ophthalmology)(National University of Malaysia) in 2000.

Her primary clinical & research interests are in glaucoma and she is currently involved in research investigating the analysis of ocular inflammatory cytokines and proteins in glaucoma, as well as novel methods for imaging, risk assessment and screening for angle closure glaucoma.

.

Email: jeanne_ogle@ttsh.com.sg