Development and Evaluation of Loop Mediated Isothermal Amplification (LAMP) for Rapid Detection of Neisseria gonorrhoeae.

SHAZIA SHAHEEN MIR, Uzair, Tanzeel, Asif and Arif Ali*

Department of Biosciences, Jamia Millia Islamia, New-Delhi, INDIA-110025

Email: shaziamir29@gmail.com, ali.arifali@gmail.com*

Nucleic acid amplification tests (NAATs) have gradually taken over the conventional methods for detection of N. gonorrhoeae (Edward T et al. 2014). However, all these NAATs have intrinsic disadvantages of requiring a specialized instrument for amplification or an elaborated and complicated method for detection of amplified products. Thus, for settings with minimal facilities, there is a need for a simple and cost-effective test that would permit rapid and reliable screening of N. gonorrhoeae (WHO, 2013). In the present study, potential of LAMP has been exploited to develop a technique that can be introduce as a “point- of-care” test in resource limited settings for the screening and identification of N. gonorrhoeae. Opa gene and porA pseudogene based LAMP assay were developed and clinically evaluated in a preselected patient population separately. A total of 388 samples were collected and evaluated, out of which 142 were male samples and 246 were female samples. Four of the 6 discrepant samples were sequenced and confirmed. Samples which were positive by Culture/or 2 gene PCRs/or sequencing were considered as “True positives” (n= 80). LAMP showed concordance (99.2%) with PCR and very high sensitivity, specificity and accuracy. With Cohan’s Kappa statistics, LAMP and PCR showed substantial agreement (99.7%) and moderate chance agreement (67.4%) with each other. Detection of the end product with the naked eyes, using SYBR Green I and Hydroxy Napthol Blue (HNB) dyes made LAMP feasible for field use also. This LAMP assay offers a highly sensitive and specific assay not only for detection, but also confirmation of N. gonorrhoeae, thereby saving cost and time.