In vitro analysis of the antibacterial effect of  the silver-coated NPWT system

Eun hee Han, researcher , Ajou University Hospital, KOREA

Bo Young Park, MD, Ewha Mockdong hospital, KOREA

Hyung Min Hahn, MD, Ajou University Hospital, KOREA

Il Jae Lee, MD, Ajou University Hospital, KOREA

Abstract (300 word limit)

Background: Post-traumatic skin and soft tissue injury is prone to infection, the risk of amputation increases when it spreads to deep tissue infection. The recently used negative pressure wound therapy (NPWT) improves the outcome of traumatic wound healing but limits the treatment of infected wounds. Therefore, the development of NPWT with anti-bacterial function is required for the care of infected skin and soft tissue wounds. Silver is a substance that has antimicrobial function without damaging human tissue. In this study, we tested the in vitro antimicrobial activity of the silver-coated NPWT system.

Methods: Bacteria strains, common pathogens of traumatic wound infections, were inoculated on agar plates. And then, the original or silver-coated foam or NPWT system was placed on bacterial inoculated agar plates. We measured the colony counts and compared the counts in the original and silver-coated foam or NPWT system for each bacteria strain, respectively.

Results: The count of colony in conventional foam was 60 for Escherichia coli, 101 for Pseudomonas aeruginosa, and 207 for Staphylococcus aureus. On the other hand, the colony count in silver foam was 18 for E. coli, 58 for P. aeruginosa, and 137 for S. aureus (Fig 1A). The colony counts were also compared in the original and silver-coated NPWT system. The colony count in the original NPWT system was 35 for E. coli, 52 for P. aeruginosa, and 81 for S. aureus. the colony count in silver-coated NPWT system was 5 for E. coli, 31 for P. aeruginosa, and 28 for S. aureus (Fig 1B).

Conclusion: In this study, the count of colonies was significantly lower than the original NPWT system in silver-coated NPWT system for all tested bacteria strains. In vivo and clinical studies on the antimicrobial effect and safety of silver-coated NPWT system are needed in the future.

Image

Figure 1. Comparison of the colony counts in conventional and silver-coated foam (A) as well as the original negative pressure wound therapy (NPWT) and silver-coated NPWT system (B).

Recent Publications (minimum 5)

- Dorafshar AH et al. A prospective randomized trial comparing subatmospheric wound therapy with a sealed gauze dressing and the standard vacuum-assisted closure device: a supplementary subgroup analysis of infected. wounds.Wounds. 2013 May;25(5):121-30.

- Stannard JP et al. Incisional negative pressure wound therapy after high-risk lower extremity fractures. J Orthop Trauma. 2012 Jan;26(1):37-42. doi: 10.1097.

- Arhi C et al. Use of a silver dressing for management of an open abdominal wound complicated by an enterocutaneous fistula-from hospital to community. J Wound Ostomy Continence Nurs. 2013 Jan-Feb;40(1):101-3.

- Matiasek J et al. The combined use of NPWT and instillation using an octenidine based wound rinsing solution: a case study. J Wound Care. 2014 Nov;23(11):590, 592-6.

- Incani V et al. Ag5IO6: novel antibiofilm activity of a silver compound with application to medical devices. Int J Antimicrob Agents. 2015 Jun;45(6):586-93.

Email: halfend@naverl.com

Notes/Comments: