Handheld ultrasound or conventional ultrasound devices in patients undergoing HSCT: A validation study

Giuseppe Milone

"G.Rodolico - San Marco" Polyclinic University Hospital, Italy

Abstract

Abdominal ultrasound exams play a major role in the diagnosis of sinusoidal obstruction Syndrome/Veno-Occlusive Disease (SOS/VOD). The development of portable Hand-Held Ultrasound Devices (HHUS) has been shown to facilitate the diagnosis of many diseases, but little data on the value of HHUS in the diagnosis of SOS/VOD are available. We performed a study aimed at validating portable Ultra-Sound (US) devices in the setting of Hematopoietic Stem Cell Transplant (HCT). Sixteen evaluable patients undergoing allogeneic HCT were studied using conventional US and HHUS during the first 3 weeks after transplant. The results obtained demonstrate that there is a close correlation between conventional and handheld ultrasound examination in the measurement of the right hepatic lobe (r=0.912, p<0.0001), the left hepatic lobe (r=0.843, p<0.0001), the Portal Vein (PV) (r=0.724, p<0.0001) and the spleen (r=0.983, p<0.0001) based on Pearson’s correlation. The same data, analyzed through Lin’s concordance correlation coefficient, evidenced a substantial level of agreement in the comparison of the spleen and right hepatic lobe, while a lower grade of agreement in the measurement of the portal vein and left hepatic lobe. Moreover, there was good agreement between results obtained by the two types of ultrasound devices in assessing ascites (p<0.0001), gallbladder thickening (p<0.0001) and the direction of PV flow (p<0.0001). HHUS device allows the study of HokUs-10 parameters with an excellent agreement with conventional US and may contribute to SOS/VOD diagnosis.

Biography

Milone Giuseppe has extensive experience in Hematopoietic transplants; he is Director of the Hematopoietic Transplantation Program in Eastern Region of Sicily (Italy). He has published over 100 peer-reviewed papers on bone marrow transplantation. He is Section Editor for the journal “Frontiers in Immunology” and for the “Journal of Clinical Medicine”. He is an opinion leader in the field of Graft versus Host disease and in the field of Hematopoietic Stem Cell Mobilization.