Establishment of a BCR-ABL p190 copy number estimator for Xpert® BCR-ABL ultra p190 test

Mengying(Mona) Liu*, Yan(Amber) Zhao, Yuanyuan Liu, Huilin Wei, Tran Tran, Lin Yuan and Gwo-Jen Day

Cepheid, United States

Abstract

Statement of the problem: BCR-ABL p190 (e1a2) is expressed in 20-30% of Acute Lymphoblastic Leukemia (ALL) and 1-2% of Chronic Myeloid Leukemia (CML). The co-expression of p190 and p210 is rare in both CML and ALL. Xpert® BCR-ABL Ultra p190, an automated cartridge-based test for measuring BCR-ABL p190 transcript level, is standardized to quantify the amount of BCR-ABL p190 relative to ABL1 control gene based on delta Ct in peripheral blood of patients. Since BCR-ABL p190 level is crucial for risk stratification and treatment decisions in ALL as well as diagnosis and continuous therapeutic monitoring in CML, it can be useful to obtain the Copy Number (CN) of BCR-ABL p190. The purpose of this study is to establish a BCR-ABL p190 CN estimator with known CN of e1a2-ABL IVT-RNA as well as to compare %BCR-ABL p190/ABL1 reporting between delta Ct-based and CN-based methods.

Methodology and theoretical orientation: Nine levels of e1a2-ABL IVT-RNA as well as two lots of Xpert® BCR-ABL Ultra p190 tests were used to generate standard curves for CN and %CN reporting. BCR-ABL p190 contrived samples and ALL clinical samples containing the BCR-ABL p190 transcript were examined to evaluate the CN and %CN between two lots of the Xpert® BCR-ABL Ultra p190 tests and to compare the delta Ct-based and CN-based methods for reporting %BCR-ABL p190/ABL1 [Figure 1].

Figure 1. Two sets of BCR-ABL Ultra p190 mutation copy number estimator for Xpert BCR-ABL Ultra p190 tests, which will provide helpful information for diagnosis and prognosis of diseases relating to BCR-ABL p190. Enter reported BCR-ABL p190 Ct and ABL1 Ct into the formulas (1) and (II) to calculate the copy number. Average copy number of BCR-ABL p190 and ABL1 from both formulas will be utilized in obtaining averaged 6BCR-ABL p190/ABL1.

Findings: Linearity was demonstrated in Ct vs. CN input for BCR-ABL p190 (R2>0.99) and ABL1 (R2>0.98). Less than 1.35-fold difference was exhibited for CN and %CN across two different lots. Less than 2-fold difference was observed in %BCR-ABL p190/ABL1 reporting between delta Ct-based and CN-based approaches.

Conclusion: A BCR-ABL p190 copy number estimator for Xpert® BCR-ABL Ultra p190 test was established, which will provide helpful information for diagnosis and prognosis of diseases relating to BCR-ABL p190.

Biography

Mengying(Mona) Liu has her expertise in molecular biology, microbiology, cell biology, gene & cell therapy, drug development and biomarker discovery. She has experience in process and assay development. Current research focuses on assay development to provide accurate and high quality IVT products for cancer diagnostic and monitoring.