## Slide 1
ABLATIVE LASER/MASS SPECTROMETRY FOR DEPTH-PROFILING AND SURFACE STUDIES – AN OVERVIEW
Presenter: Avin Pillay, The Petroleum Institute, Abu Dhabi, UAE

## Slide 2: OUTLINE
Introduction
Instrumentation/Sample Handling
Results/Discussion
Future Work

## Slide 3: INTRODUCTION
Laser Ablation-ICP-MS is a hyphenated technique that has the following advantages:
Quick
Accurate
Multi-elemental
Easy sample handling

## Slide 4: Methodology
Instrumentation
Sample Handling

## Slide 5: Ablative Laser Technique
Useful for surface uniformity studies
Has depth-profiling capabilities
Semi-quantitative
Can be used on soft sample after pre-treatment with liquid nitrogen

## Slide 6: Laser Ablation ICP-MS
Samples are irradiated with a 213 nm laser micro-beam.
The vapor is transported to ICP-MS for analysis

## Slide 7: Atomization/Ionisation in Plasma
Vapors are atomized and ionized in hot plasma and the ions are collimated

## Slide 8: Quadrupole mass selection for fine resolution
A set four magnets deflects ions of varying mass to the mass detector

## Slide 9: Depth-Profiling
The laser “drills” through the sample and measures the characteristic intensity of metals beneath the surface

## Slide 10: Surface studies – Grid Analysis
The laser scans the surface at selected points on a grid
Uniformity of signal obtained predicts surface uniformity

## Slide 11: Validation of ICP-MS
Technique is calibrated with certified standards in the aqueous phase to establish performance of the system
Element | Certified value (ppb) | Measurement 1 Fluka 70007 (ppb) | Relative Error | Measurement 2 Fluka 70007 (ppb) | Relative Error
Be | 10 | 10.20 | +2.0% | 10.75 | +7.5%
Mg | 10 | 9.32 | -6.8% | 9.92 | -0.80%
Co | 10 | 9.74 | -2.6% | 9.98 | -0.2%
Ni | 10 | 9.93 | -0.70% | 10.10 | +1.0%
In | 10 | 9.94 | -0.60% | 9.94 | -0.60%
Pb | 10 | 10.70 | +7.0% | 11.19 | +11.2%
Bi | 10 | 10.06 | +0.60% | 10.36 | +3.6%

## Slide 12: Results/Discussion
Impact of the study

## Slide 13: Homogeneity in Doped Polymers
General uniformity of dopant in sample can be established by studying consistency of signal output
Aberrations in beam strength could create slight fluctuations

## Slide 14: Untreated Asphaltenes
Some asphaltene samples are sufficiently coalesced to withstand direct irradiation without splashing effects

## Slide 15: Metal inclusions/Hotspots in untreated asphaltenes
Nickel hotspots observed at different depths
Evidence of isolated metal inclusions

## Slide 16: Screen-shot of “splashing” effects in a soft sample
If samples are soft and gelatinous, splashing occurs on irradiation with the laser as shown below

## Slide 17: Spectrum of a “splashed” gelatinous sample showing erratic features

## Slide 18: No splashing indicates sample is solid

## Slide 19: Treatment of samples with liquid nitrogen
Samples are solidified by pouring liquid nitrogen over them – see wax sample below
The solid sample is then quickly subjected to laser ablation

## Slide 20: Point of thawing in sample chamber
Analysis must be done before thawing
Thaw time can be roughly 3 minutes

## Slide 21: Typical depth-profiling spectra from samples pre-treated with liquid nitrogen
Spectrum of a petrified wax sample shows from consistency of peaks that the liquid nitrogen approach is successful

## Slide 22: Future Studies: Carbon in steels
Our next study is to measure homogeneity of carbon in steels to determine integrity and uniformity
This can be extended to evaluation of nickel, manganese, chromium and cobalt in steels

## Slide 23: ACKNOWLEDGEMENTS
The Petroleum Institute for financial support
Mr. Sasi Stephen for operating the laser ablation system
Collaborators for their assistance

## Slide 24: THANK YOU
QUESTIONS