## Slide 1: Comprehensive Two-Dimensional Liquid Chromatography for Effect-Directed Analysis in Effluents and Surface Water
Xiyu Ouyang, Jacob de Boer

## Slide 2: Objective
Develop an LCxLC system linked to a fraction collector and
high resolution mass spectrometer to support the identification
and monitoring of emerging toxicants in effect-directed
analysis

## Slide 3: Instrumentation: LCxLC-ToF MS/UV-FC
pump 1 (HPLC)
pump 2 (UPLC)
auto-sampler
ESI-TOF-MS
UV detector
fraction collector
waste
1D column
2D column
loop 1
loop 2
2DLC interface
20%
80%
pump 1 (HPLC)
pump 2 (UPLC)
auto-sampler
ESI-TOF-MS
UV detector
fraction collector
waste
1D column
2D column
loop 1
loop 2
2DLC interface
20%
80%

## Slide 4: LC×LC-TOF MS Analysis
Dissect: S/N>15
More than 700 compound spectra
Manual “subtraction” necessary
Peak intensity < 10000 counts
Peak intensity > 80000 counts

## Slide 5
Mode | IEC× RP | SEC×RP | NP× RP | RP× RP | HILIC× RP | HILIC× HILIC | SEC×NP | SEC×IEC
Orthogonality | ++ | ++ | ++ | + | + | - | + | +
Peak capacity | + | + | + | ++ | + | + | - | --
Peak Capacity/Time | - | -- | + | ++ | + | + | -- | --
Solvent compatibility | + | + | -- | ++ | + | ++ | + | +
Applicability | + | + | - | ++ | + | - | - | -
Score | 4 | 3 | 1 | 9 | 5 | 2 | -2 | -3
Xiaoping Li et al. Journal of Chromatography A, 1218 (2011) 2214–2221
Stationary phases selection

## Slide 6
Three different stationary phases in the 2nd dimension (1st dimension C18) for comparison:
A: Coreshell PH 2.7µm 50x4.6 mm
B: Phenomenex Kinetex PFP 2.6µm 50x4.6 mm
C: HILIC 1.8µm 50x3.0 mm
1D C18 0.1ml/min 60min; 2D 2ml/min 30s H2O/ACN, UV at 290nm
Orthogonality estimation:
Surface coverage
(vector cross product)
A: 0.643
B: 0.671
C: 0.687
Conclusion: C18×PFP
Stationary phases comparison

## Slide 7: EPA 8270c Mixture (38 compounds)
Semi-volatile compounds from solid waste, soil, air, and water
Acenaphthene | 2,4-Dimethylphenol
Acenaphthylene | Fluoranthene
Acetophenone | Fluorene
Aniline | Hydroquinone
Anthracene | Indeno[1,2,3-cd]pyrene
Benzoic acid | Isophorone
Benz[a]anthracene | 3-Methylcholanthrene
Benzo[b]fluoranthene | 2-Methylnaphthalene
Benzo[k]fluoranthene | 2-Methylphenol
Benzo[ghi]perylene | 3-Methylphenol
Benzo[a]pyrene | 4-Methylphenol
p-Benzoquinone | Naphthalene
Benzyl alcohol | Perylene
Chrysene | Phenanthrene
Dibenz[a]acridine | Phenol
Dibenz[a,h]anthracene | 2-Picoline
Dibenzofuran | Pyrene
Dibenzo[a,e]pyrene | Pyridine
7,12-Dimethylbenz[a]anthracene | Resorcinol
ACN/water system, 30oC, UV detection at 290nm

## Slide 8: Triazine and Urea Pesticide Mixture (29 compounds)
Atrazine | Isoproturon
Atrazine-desethyl | Metabromuron
Atrazine-desisopropyl | Metazachlor
Metamitron | Propazine
Chloridazon | Dimefuron
Metoxuron | Terbuthylazin
Carbetamid | Linuron
Bromacil | Chloroxuron
Simazine | Prometryn
Cyanazin | Chlorpropham
Terbutylazine-desethyl | Terbutryn
Methabenzthiazuron | Metolachlor
Chlorotoluron | Ethofumesate
Monolinuron | Ethidimuron
Diuron | 
ACN/water system
Ambient temperature
ESI-TOF detection, 27/29 detectable

## Slide 9
Reproducibility (2D retention time stability)
(A) methomyl
(B) 3-hydroxycarbofuran
(C) propoxur
(D) carbofuran
(E) carbaryl
(F) methiocarb

## Slide 10: LC/MS Conditions
ToF-MS resolution 10,000
ESI: 4500 V with end plate offset -500 V
Flow rate 2nd dimension: 400 µl/min
Agilent 1260 fraction collector, delay volume: 55 µl
Splitter 1:4
m/z 50-1,000
Nebulizer gas N2, 4.0 bar
Drying gas: 8 L/min, 200 ˚C
Capillary exit was set at 100 V, skimmer voltage 33.3 V
Hexapole RF: 90 Vpp and lens 1 pre-pulse storage: 1 µs
Bruker OtofControl 3.0 (Bruker Daltonics), 5 Hz scan

## Slide 11
Example: effluent sample from WWTP Cuijk
Courtesy BE-Basic project
6 sampling location in NL (surface water, effluent)
silicon rubber passive sampler
large volume solid phase extraction

## Slide 12: Silicon rubber treatment
Silicon rubber blades were cut into small pieces and
collected in pre-cleaned thimbles in the Tecator® Soxtec Avanti
The extraction with 80 ml of a methanol: acetonitrile (1:2 v/v)

## Slide 13
LC×LC of Cuijk WWTP sample
1
2
3
4
Diazinon
Terbutryn

## Slide 14: Contour plot

## Slide 15: Tentative identification
No. | m/z measured | Intensity | Suggested molecular formula [M+H]+ or [M+Na]+ | m/z calculated | error (ppm) | mSigma | Tentative candidates |  | Application/Source
4 | 202.1225 | 9828 | C13H16NO | 202.1226 | 0.51 | 7.3 | (E)-1-Cinnamoylpyrrolidine | food and beverage additive | 
5 | 156.0815 | 6944 | C11H10N | 156.0809 | -0.69 | -4.5 | 4-Phenylpyridine | beverage additive | 
6 | 278.1900 | 6779 | C20H24N | 278.1903 | 1.30 | 10.1 | Amitriptyline or EDDP | antidepressant/ drug metabolite | 
7 | 207.1493 | 9287 | C12H19N2O | 207.1492 | 0.24 | 8.3 | Isoproturon | herbicide | 
12 | 212.2010 | 8166 | C13H26NO | 212.2009 | -0.08 | 9.1 | Ethyl menthane carboxamide | food and beverage additive | 
14 | 293.1055 | 10882 | C13H19ClN2NaO2 | 293.1027 | -9.40 | 12.6 | Chloroprocaine | local anesthetic | 
15 | 182.0095 | 7237 | C8H8NS2 | 182.0093 | -1.35 | 14.1 | 2-Methylthiobenzothiazole | degradation product of biocides | 
17 | 232.1334 | 33911 | C14H18NO2 | 232.1332 | 0.92 | 0.7 | Indeloxazine | drug for cerebrovascular disease | 
18 | 242.1427 | 4947 | C10H20N5S | 242.1434 | -1.76 | 6.7 | Terbutryn | herbicide | 
20 | 225.0914 | 8743 | C15H13O2 | 225.0910 | 1.57 | 3.3 | Flavanone | natural product | 
22 | 226.1337 | 11064 | C14H16N3 | 226.1339 | 1.00 | 11.6 | Cyprodinil | pesticide | 
23 | 219.1945 | 16676 | C12H27O3 | 219.1955 | -4.55 | 8.3 | 8,8-Dimethoxy-2,6-dimethyl-2-octanol | food additive | 
 | 241.1773 | 10252 | C12H26NaO3 | 241.1774 | -0.70 | 3.2 | Na adduct | / | 
26 | 312.2317 | 7057 | C21H30NO | 312.2322 | -1.66 | 9.5 | Biperiden | drug (Parkinson’s disease) | 
27 | 305.1094 | 34662 | C12H22N2O3PS | 305.1088 | -2.29 | 5.9 | Diazinon | pesticide | 
28 | 322.0665 | 29395 | C16H17ClNO2S | 322.0663 | -0.64 | 2.8 | Clopidogrel | drug for cerebrovascular disease | 
29 | 399.2510 | 159335 | C18H40O7P | 399.2506 | 0.96 | 4.0 | Tris(2-butoxyethyl) phosphate | organic flame retardant | 
 | 421.2326 | 24812 | C18H39NaO7P | 421.2326 | 0.02 | 2.4 | Na adduct | / | 

## Slide 16
Surface water sample from Lekkanaal
Samples from BE-Basic project, courtesy to all relevant

## Slide 17
LCxLC Chromatogram of water sample Lekkanaal

## Slide 18
Biological
analysis
Fractionation
Chemical
analysis
List of
toxicants
List of
suspects
Extract
Confirmation
Identification
Effect-directed Analysis (EDA)
  Demand of developing/applying a technique that offers greater separation power: LC×LC!

## Slide 19
Zebrafish embryo toxicity test (zFET)
Embryo (2 days)
Larva (3 days)
Freshwater fish, native to Southeastern Himalaya region
Egg-diameter is around 1 mm, survives in 96 well plates till day 6
Clear chorion allows to monitor the early development
Till the end of the larva stage (day 7) they are not “animals” (replace animal tests)
Complex organism - > Multiple Endpoints
1 hour

## Slide 20: Examples of embryotoxicity
control
thiram
bAPN
RA

## Slide 21: Zebra Fish: Model Organism for Teratogenicity Studies

## Slide 22
EDA Study: LC×LC-ZFET
LC x LC
Fractionation
Phenotype screen
Chemical-Mix
Environmental Samples
LC/MS
Chemical Identification
12h on shaker
Centrivap 6h
Add Fish-medium
Add Zebrafish eggs

## Slide 23: Summary
LC×LC (ToF-MS) is a very promising tool in analyzing complex environmental samples due to greater peak capacity and multi-(tunable) selectivity.
Direct post-column high resolution fractionation of LC×LC combined with high throughput bioassay makes EDA easier and better.
Future plan:
Application of LC×LC facilitated EDA in combination of a high-throughput fluorescent TTR assay to screen for Thyroid hormone disruptor of indoor environment, where:
  Novel sample preparation method will be developed
  More stationary combinations will be tested
  Different ionization sources (APCI, APPI, etc.) will be applied
FI Enhancement

## Slide 24: Acknowledgement
Dr. Jessica Legradi
Dr. Ron van der Oost
Prof. Dr. Juliette Legler
Dr. Marja Lamoree
EU project (FP7, Marie Curie ITN) EDA EMERGE
X. Ouyang et al. (2015), J. Chromatogr. A 1380, 139-145