## Slide 1: About OMICS Group
OMICS Group is an amalgamation of Open Access Publications and worldwide international science conferences and events. Established in the year 2007 with the sole aim of making the information on Sciences and technology ‘Open Access’, OMICS Group publishes 700+ online open access scholarly journals in all aspects of Science, Engineering, Management and Technology journals. OMICS Group has been instrumental in taking the knowledge on Science & technology to the doorsteps of ordinary men and women. Research Scholars, Students, Libraries, Educational Institutions, Research centers and the industry are main stakeholders that benefitted greatly from this knowledge dissemination. OMICS Group also organizes 1000+ International conferences annually across the globe, where knowledge transfer takes place through debates, round table discussions, poster presentations, workshops, symposia and exhibitions.

## Slide 2: OMICS International Conferences
OMICS International is a pioneer and leading science event organizer, which publishes around 700+ open access journals and conducts over 500 Medical, Clinical, Engineering, Life Sciences, Pharma scientific conferences all over the globe annually with the support of more than 1000 scientific associations and 30,000 editorial board members and 3.5 million followers to its credit.
OMICS Group has organized 1000+ conferences, workshops and national symposiums across the major cities including San Francisco, Las Vegas, San Antonio, Omaha, Orlando, Raleigh, Santa Clara, Chicago, Philadelphia, Baltimore, United Kingdom, Valencia, Dubai, Beijing, Hyderabad, Bengaluru and Mumbai.

## Slide 3
Jim Polarine MA.
August 17, 2015
Case Studies of Human Flora and Spore Contamination in Cleanrooms

## Slide 4
Case Studies in Bioburden Control
Starting Up a New Cleanroom
Agenda

## Slide 5: Bioburden Agenda
Operator Contamination
Fungal Spore Contamination
Bacterial Spore Contamination

## Slide 6: Operator contamination
Staphylococcus
Propionibacterium acnes

## Slide 7: Human Skin Flora
Staphylococcus cleanroom outbreak
  Source
    Gowning Material
    Traced back to one operator
    Skin infection
    Another case was from a Tongue Stud
    Non-sterile drug product
Courtesy Ann Larson

## Slide 8: Propionibacterium acnes
Gowning room
  BSL Hood
  Sources
    Operator with acne
      Solution
        Antimicrobial Handwash
        Antimicrobial Bodywash
Courtesy Ann Larson

## Slide 9: Agenda
Operator Contamination
Fungal Spore Contamination
Bacterial Spore Contamination

## Slide 10: VALIDATION – Microorganism Selection
From McDonnell, “Antisepsis, Disinfection, and Sterilization: Types, Action, and Resistance” 2007, ASM Press

## Slide 11: Fungal Spores
Penicillium
Aspergillus
Cladosporium

## Slide 12: Common sources of Spores
Items brought into the Cleanroom
  Bags, Boxes, Intervention Equipment, Pallets, Pallet Jacks, Scrubbers, Cart Wheels, Shoes, Shoe Covers
  Raw Materials

## Slide 13: Penicillium
ISO-7 Cleanrooms
Action Levels of 10 and picking up >100
  Engineering Investigating
  HVAC
  Duct Work
  HEPA Filters
  Cooling Coils
  Wall Coverings
  Airflow Vents

## Slide 14: Penicillium Investigation
Entry and Exit Procedures
Gowning Procedure
Cart Wheels
Construction
  Further Investigation
  Use of Sporicides
  Containers in the Cleanroom
  Coldroom Cleaning Procedures
  Documentation
  Assignable Cause

## Slide 15: Penicillium
Courtesy Ann Larson

## Slide 16: Aspergillus
ISO-5 Cleanroom
  Source
    Door Kick Plate
    High Impingement Spraying Device
Exceeding Limits in ISO-7 areas
  Dock Doors proximal to ISO-7 cleanroom
  Storage room with limited control
  No limits for mold spores
  Limited control for incoming and outgoing items

## Slide 17: Aspergillus Investigation
Sporicidal usage in pass thru
Clean and dirty area on the Dock
Better control of gowning area
Cart Wheel control
Better gowning control

## Slide 18
Fungi
Courtesy Dan Klein

## Slide 19
Cleanroom Fungi
Courtesy Dan Klein

## Slide 20: Inoculum Preparation—Fungal Spores
Cultures need to be incubated for a sufficient length of time before harvesting spores
Courtesy Dan Klein

## Slide 21: Surface Preparation
Autoclaving may not be acceptable for some surfaces (Saniflex)
Courtesy Carol Bartnett

## Slide 22: Surface Tension Issue

## Slide 23: Cladosporium
Elevated levels in Puerto Rico
Common in Southern CA
  Increase Sporicide Usage
  Sources
    Humidity and temperature
    Burning of Sugar Cane fields

## Slide 24: Case Study on Substrates
Efficacy (log reduction) of Low pH phenolic (Environ LpH se) : (1:256 Dilution) against test microorganisms on representative surfaces |  |  |  |  |  | 
Surface | Staphylococcus epidermidis | Pseudomonas aeruginosa | Corynebacterium glutamicum | Candida albicans | Aspergillus niger | Penicillium chrysogenum
Stainless Steel | 6.62 | >6.10 b | 4.18 | >4.31 b | <3.00c | 4.95
Glass | 6.85 | 6.42 | 5.26 | >5.80 b | 2.98 | 5.11
Aluminum | 6.35 | 5.69 | 5.14 | >3.93 b | <3.00c | 3.48
Epoxy | 4.36 | 4.45 | 4.48 | 3.19 | <3.00c | <3.00c
Enamel | >6.05b | >5.72 b | 5.45 | >3.92 b | <3.00c | 2.83
Acrylic | 4.53 | 6.06 | 4.49 | 2.92 | <3.00c | <3.0 c
Mipolam | 4.36 | 3.87 | 4.29 | 4.37 | <3.00c | 3.25
Vinyl | 4.08 | 3.68 | 3.93 | 2.61 | <3.00c | 2.1
Hardwood | 5.18 | >4.54 b | 5.26 | 3.2 | <3.00c | 2.59
Melamine Covered Wood | >5.38 b | >5.64 b | >5.09 b | >5.12 b | 3.65 | 3.95
Plastic | >5.73 b | >5.32 b | >5.05 b | >4.04 b | <3.00c | 2.44
Plexiglas | >5.90 b | 5.62 | 4.83 | >4.40 b | <3.00c | 3.85
Print | 5.85 | 5.86 | 5.74 | 4.51 | <3.00c | 3.38
Chromium | 6.55 | 5.95 | 6.63 | 4.08 | <3.00c | 2.61
a Disinfectant Efficacy = (Log MSP(positive control) - Log MSP(test coupons)), where MSP(Positive Control)= Mean surviving population on positive control coupons; MSP(test coupon) = Mean surviving population on test coupons after disinfectant treatment; b Each of triplicate coupons showed no growth after disinfectant treatment; c Each of triplicate coupons showed TNTC growth |  |  |  |  |  | 

## Slide 25: 70% IPA Efficacy Against Molds

## Slide 26: H2O2/PAA RTU Against Molds

## Slide 27: Agenda
Operator Contamination
Fungal Spore Contamination
Bacterial Spore Contamination

## Slide 28: Bacterial Spores
Bacillus cereus
Bacillus circulans

## Slide 29
Bacterial Endospore
Courtesy Dan Klein

## Slide 30
Bacterial Endospore
Bacterial Spore Structure

## Slide 31: Bacillus cereus
ISO-7 and ISO-8 cleanrooms
Process Vessels
  Source Locations
    Cleanroom Shoe Cover
    Fermentor
    Process Vessels
    The Source was a Raw Material

## Slide 32: Exosporium – B. anthracis
Cote CK et al. 2011. Microbes and Infection 13(14-15):1146-55.

## Slide 33: Exosporium – B. anthracis
https://www.llnl.gov/str/Sep06/Velsko.html
Hydrophobicity helps adhere to fibers

## Slide 34: Bacillus Testing
H202/PAA 1:18
Aklaline Cleaner-2
Iodine Product
Acidic Cleaner
Alkaline Cleaner
1:10 Bleach
1:4 Bleach
H202/PAA RTU

## Slide 35: Bacillus circulans
Efficacy Testing Failures (different test methods)
  Vinyl surfaces
    Sent out for ID
      Paenibacillus

## Slide 36: Case Study: Construction Event at Biotech Site
Worst Case Events
9X Clean [1X Sporicide + 2X Phenolic (day 1) 2X Phenolic +1X Sporicide (day 2), 2X Phenolic +1X Sporicide (day 3)]
Fogging
VHP
Triple Clean
  Defined 3X Disinfectants and Sporicide
  EM frequency (Static and Dynamic)
  Release of the room

## Slide 37: Cleaning and Disinfection Efficacy

## Slide 38: Time 0
Red = Sporeformers Green = Other

## Slide 39: After 1X Cleaning - NO Sporicide

## Slide 40: After 2X Cleaning – NO Sporicide

## Slide 41: After 3X Cleaning - No Sporicide

## Slide 42: After Sporicide

## Slide 43: Bioburden Agenda
Operator Contamination
Fungal Spore Contamination
Bacterial Spore Contamination

## Slide 44
Summary
Case Studies in Bioburden Control
Starting up a New Cleanroom

## Slide 45
Thank You for Your Time
Jim Polarine Jr. MA.
jim_polarine@steris.com

## Slide 46: Let us meet again..
We welcome you all to our future conferences of OMICS International
2nd International Conference and Expo on
Parenterals and Injectables
On
October 24-26, 2016 at Istanbul, Turkey
http://parenterals-injectables.pharmaceuticalconferences.com/