## Slide 1
Method for Hot Real-Time Sampling of Gasification Products
Bioenergy 2016 – Rome, Italy
Marc Pomeroy
June 14, 2016

## Slide 2: Thermochemical Biomass Conversions
Rapid heating of biomass to around 500°C in absence of oxygen
  Primary product is complex bio-oil
    Over 400 compounds
    Chemistry changes when condensing from vapor to oil
    Acidic
    High oxygen content
    Up to 60% of material does not re-volatilize making analysis challenging
Rapid heating of biomass to around 825°C in presence of controlled oxygen or steam
  Primary product Syngas (H2 + CO) for liquid fuel synthesis
    Still a complex mixture with tars and inorganic contaminants
    Up to 70% steam content
    Formation of catalytic poisons such as hydrogen sulfide

## Slide 3: Finding Solutions for Biomass Conversion at NREL
Feedstocks
Scale-up & Demonstration
100’s kg
Technoeconomic Analysis
Bio-Oil Characterization
Fundamental Science
mg-g
Catalyst Development & Testing
g-kg
Research at multiple scales from fundamental, to bench, to pilot plant.
Overarching research necessary to support lab and industrial deployment.
Gasification Products

## Slide 4: Analytical Procedures – Validated Methods
WWW.NREL.GOV/BIOENERGY

## Slide 5: Thermochemical platform at NREL has multiple systems from mg to 450kg/day scales and operating in a variety of configurations
Small Scale Reactors: Catalyst Development
Catalyst use per test: 0-2g
Laminar Entrained Flow Reactor: In-Situ Pyrolysis
Biomass rate: <5 g/hr
2” Fluidized Bed Reactor:
Fast, Ex-situ, & In-Situ Pyrolysis
Biomass rate: <0.5 kg/hr
4” Fluidized Bed Reactor: Gasification
Biomass rate: <2 kg/hr
Davison Circulating Riser: Ex-Situ Pyrolysis
Biomass rate: <5 kg/hr
Thermochemical Process Development Unit: All Pathways
Biomass rate: <30 kg/hr

## Slide 6: Pilot Scale Research Configurations
The Thermochemical Process Development Unit (TCPDU) is designed to be flexible and operate in three primary configurations:
  Fast Pyrolysis
    Production of condensed Bio-Oils
  Ex-Situ Catalytic Fast Pyrolysis
    Catalytic upgrading of Pyrolysis vapors prior to condensation
  Gasification
    Production and upgrading of Syngas

## Slide 7: Process Flow Diagram – TCPDU
Hot Sample Port
Hot Sample Port
Condensed
Sample

## Slide 8: All systems are capable of hot real-time sampling
Molecular Beam Mass Spectrometer (MBMS)
  Sampling up to 500° C
  Supersonic expansion, rapid cooling/rarefaction preserves sample without condensation or reaction
  Mass analysis provides instantaneous chemical fingerprint of on-line sample

## Slide 9: MBMS Representative Tar Spectrum
m/z

## Slide 10: TCPDU – Modified Agilent GCs
Hot sampling Gas Chromatographs (GC)
  Valve material limited to 325°C
  Sulfur Chemiluminescent Detector (SCD)
    Trace Measurements to 0.05 ppm
  Nitrogen Chemiluminescent Detector (NCD)
    Trace Measurements to 0.01 ppm

## Slide 11: Challenges of Sampling at the Pilot Scale
Analytical equipment adversely affected by high ambient temperatures/dust
Space around reactors required for plant maintenance and operation
Multiple sampling points
  Pre- and Post-catalytic Reactors
Plugging and flow issues
Residence time
  Thermal changes
    Cracking, further reaction chemistries
Catalytic effects
  Ash, char, anti-seize
Material Selection
  Sulfur interactions with steel
Operational Constraints
Representative Sample

## Slide 12: Guide to Successful Sampling – Sample Ports
Position orthogonal prior to bends
  Avoid in-line placement
Place on upper half of pipe to prevent accumulated condensates from entering sample lines
Process Stream

## Slide 13: Guide to Successful Sampling - Filtration
Filter as close to process as possible
  Filter volume large enough to prevent particulate plugging
  Sample volume small enough to prevent dilution or long residence times
    Cracking of samples or catalytic changes due to particulate interaction
  Appropriate filtration size for sampling equipment and tubing
    TCPDU uses 10-15 μm elements
  Materials/coatings selection
    Sulfur interactions

## Slide 14: Guide to Successful Sampling – Materials Selection
Sulfur species tend to be Catalyst Poisons
  Higher concentrations from Herbaceous Feedstocks – Primarily Hydrogen Sulfide (H2S) in Syngas
    400 ppm to 800 ppm H2S from Corn Stover
    5 ppm to 50 ppm H2S from Oak
Hydrogen Sulfide binds to steel surfaces at temperature and can take significant time to reach steady state at low concentrations
  Use Stainless Steels (316L)
  Coat with commercially available coatings
    Keep Below 500°C
    Do NOT “Burn out” lines or allow oxygen at temperature

## Slide 15: Effect of Gasket Material - Change of Sulfur Species
Chart: Sulfur in ppm
COS in ppm
Chart: Sulfur in ppm
3" coated Housing, coated Pleated Filter, Graph-Lock Gasket
Notes: With all coatings combinations and the use of a Graph-Lock Graphoil gasket, conversion of H2S (Hydrogen Sulfide) to COS (Carbonyl Sulfide) was observed in significant amounts. Use of a ceramic fiber gasket did not show the same COS formation, but was very difficult to achieve a gastight seal.

## Slide 16: Materials Limitations - Valves
High Temperature multi-position valves limited to 325°C

## Slide 17: Guide to Successful Sampling – Even Heating
Sample line temperatures must not change composition of Gas Stream
  Gasification much more forgiving than pyrolysis
    Stay between 425°C – 500°C
  Be careful of heat sinks like fittings
    Can partially or fully condensate a sample and promote plugging

## Slide 18: Gasification products are hard on analytical equipment
Robust pumping systems required
Coatings limited to <500°C
Cannot “Burn Out” coated components
Heated nitrogen purge when not actively sampling
Redundant sample lines
Material compatibility
  Valve packing materials my pyrolize at temperature

## Slide 19: Robust Sampling System - Summary
Leave time for commissioning of sampling system
  May need to optimize for an individual feedstock
Place your sample ports effectively
Filter close to the process
Carefully select materials and coatings
  Do not oxidize or exceed 500°C
Even heating of sample components is critical
Solvent cleaning of residual tars may be required
Analytical equipment must be designed for high temperatures and condensable components
Adequate sample flow for as short a sample time as reasonable

## Slide 20: Acknowledgements
TCPDU Team
Dr. Esther Wilcox - PI
Daniel Dupuis
Tim Dunning
Katherine Gaston
Chris Golubieski
Raymond Hansen
Matthew Oliver
Kristin Smith
Other Contributors
Daniel Carpenter
Steve Deutch
Tina Fehringer
Calvin Feik
Dr. Mark Jarvis
Marc Oddo
Jack Ferrell
US Department of Energy
Bioenergy Technologies Office

## Slide 21
Questions?