## Slide 1: Metallic 3D-printing for orthopedic surgery: question of surface and cell compatibility
Michaela Fousová, Dalibor Vojtěch, Eva Jablonská, Jaroslav Fojt

## Slide 2
Introduction
Metallic 3D printing principle of powder-bed techniques
Experimental setup SLM and EBM technology
Results surface morphology and chemistry, cell compatibility
Conlusions
Outline

## Slide 3
University of Chemistry and Technology Prague, Czech Republic
Department of Metals and Corrosion Engineering
Introduction

## Slide 4
3 main categories of manufacturing process:
Powder bed systems
  raking powder across the building plate to form successive layers
  an energy source (electron or laser beam) selectively melts the powder to form desired shape
  advantage: high resolution, dimensional control
Powder feed systems
  powders are conveyed through a nozzle onto the building plate
  a laser is used to melt a monolayer or more of the powder into the shape desired
  advantage: larger build volume, reparation of damaged components
Wire feed systems
  feed stock is wire which is deposited onto the building plate
  upon subsequent passes a 3D structure is built up
  welding by laser, electron beam or plasma arc
  Advantage: high deposition rate
Metallic 3D printing

## Slide 5
Powder bed systems
Selective Laser Melting (SLM) and Electron Beam Melting (EBM)
Metallic 3D printing

## Slide 6
SLM
laser beam, 200 W
argon atmosphere
Input Ti6Al4V powder:
15-45 μm
Layer thickness: 30 µm
no preheating
heat treatment necessary
EBM
electron beam, 60 kV
vacuum
Input Ti6Al4V powder:
45-105 µm
Layer thickness: 50 µm
preheating every layer at 740 °C
heat treatment not necessary
Metallic 3D printing

## Slide 7
SLM
ProSpon spol. s r.o., Czech Republic
development, manufacture and distribution of medical implants and instruments for orthopedics, traumatology and surgery
EBM
Metal Industries Research & Development centre, Taiwan
research & development of the leading technology of metal and its related industries in Taiwan
Metallic 3D printing

## Slide 8
high mechanical properties
excellent corrosion resistance
biocompatibility
much higher stiffness over bone
bioinert - no biological fixation
Ti6Al4V alloy

## Slide 9
SLM machine M2 Cusing, ConceptLaser
200W Yb:YAG fiber laser
protective argon atmosphere
Input material: gas-atomized Ti-6Al-4V powder
        (rematitan® CL, Dentaurum, 15-45 μm)
Post-production heat treatment: 820 °C, 1.5 h
S400 Arcam machine
electron beam, 60 kV
vacuum
Input material: gas-atomized Ti-6Al-4V powder
        (Arcam, 45-105 μm)
Experimental

## Slide 10
Process parameters:
SLM – island scanning strategy, EBM – continuous scanning
Specimens:
blocks of which specimens of 8x8x3 mm3 were cut off
Experimental
 | Power (W) | Scanning speed (mm/s) | Hatching distance (μm) | Layer thickness (μm)
SLM | 200 | 1250 | 80 | 30
EBM | 900 | 4530 | 200 | 50

## Slide 11
Characterization of surface
Morphology – SEM (Tescan Vega-3 LMU)
Analysis of chemical composition – XPS (ESCAprobe P, Omicron Nanotechnology Ltd.)
Experimental

## Slide 12
In vitro contact tests
U-2 OS cells, seeding density 24.000 cells/cm2
DAPI/phalloidin-TRITC staining (Sigma)
LIVE/DEAD® Viability/Cytotoxicity Assay (Thermo Scientific)
fixation for SEM observation
Fluorescence microscopy – fluorescence confocal microscope Olympus IX81
Experimental

## Slide 13
Results
SLM
EBM
Surface morphology
av. particle size ~ 75 µm
→ higher roughness
Adhering powder particles
av. particle size ~ 30 µm
→lower roughness

## Slide 14
Results
Surface morphology
Arcs of solidified melt

## Slide 15
Results
Significant changes
SLM
  2fold increase in Al
  no V
  except of Ti and Al oxides, also Ti carbide detected
  Cause: residual oxygen in the argon atmosphere, high temperature of laser beam
  termodynamical explanation
  EBM
  increase in Al, but also in V
  Cause: moisture in powder-bed
  TiO, TiO2, Ti2O3, but also elemental Ti → oxidic layer thinner
Surface chemistry
Chart: wt. %
| Ti | Al | V
commercial production | 84.5 | 13.6 | 1.9
SLM | 77.8 | 22.2 | 0
EBM | 78.6 | 17.3 | 4.1
Biocompatibility preserved
  Al in the form of inert Al2O3
  V in the form of V2O5

## Slide 16
Results
24 hours
72 hours
Cytocompatibility
SLM
EBM
control

## Slide 17
Results
Only comparative chart
Difficulty to assess real number of cells due to high roughness of 3D-printed surfaces
87 and 96 % of positive control (cell culture dish) for SLM and EBM, respectively
60 a 74 % of commercially produced Ti6Al4V with smooth surface
Cytocompatibility
Cell density after 24 hours of cultivation
After 1st day, cells grew preferentially on smooth surface in between adhering powder particles representing too large obstacles. Later, whole surface overgrown.
Chart: cell density (103 cm-2)
| SLM | EBM
3D-printed | 26 | 23
cell culture dish | 30 | 24
commercialy produced | 44 | 31

## Slide 18
Results
EBM
Cytocompatibility
control
healthy cell morphology

## Slide 19
Results
EBM
Cytocompatibility
SLM
control
72 hours of cultivation

## Slide 20
Results
SLM
EBM
Cytocompatibility
preferentially on smooth melted surface
SEM images of cells
24 hours of cultivation
on tested materials
control
small adhering powder particles already overgrown
healthy polyhedral shape

## Slide 21
Powder bed technologies finish in product with surface covered with adhering powder particles.
EBM yields surface of higher roughness than SLM.
Surface chemistry is changed by 3D printing process.
Biocompatibility of Ti6Al4V is retained.
Difficulty to simply assess cytocompatibility quantitatively – need of advanced techniques
Risk of particles release – need of surface treatment
3D-printed surface is suitable for orthopaedic applications, but needs to be improved.
Conclusions

## Slide 22: Thank you for your attention!