## Slide 1: Brief comparison of LC-esi-MS and LC-APCI-MS
Wenjie Cao
SABIC Technology Center
Riyadh, Saudi Arabia

## Slide 2
1st difference
2nd
3rd
Courtesy of Shimazu.com
Less background in APCI
Heater in APCI
Proton Exchange and Ion “Evaporation”

## Slide 3
Courtesy of LGC.com
A closer look of apci Ionization

## Slide 4
Pros and Cons of ESI and APCI |  |  |  |  | 
ESI |  |  | APCI |  | 
ESI Ionization | Pros | Cons | APCI Ionization | Pros | Cons
"Evaporate" ions out of droplets | Multiple charged ions for high MW chemicals | Could be difficult in MS interpretation | Gas Phase Proton Exchange | Easier to get M+1 for identification | Less sensitive or can't handle ionic chemicals
 | Metal ion for confirmation of M? | Metal adducts |  | No multiple-charged ions, easier for identification | 
 |  | May mess up with MRM |  | No metal ions, easier for identification | 
 |  | More background ions |  | Easier for MRM | 
 |  | Ion suppression |  | Neater background | 
Lower Temperature | Good for thermally unstable chemicals | More adducts and clusters | Higher Temperature | Less cluster/adduct ions | Thermally unstable chemicals?
 |  | Can't handle higher flow rate |  | Can ionize less polar chemicals | 
 |  |  |  | Can handle higher LC-flow rate | 
Brief Comparison of ESI and APCI

## Slide 5
LC FLOW RATE FOR ESI AND APCI
Courtesy of F. Klink

## Slide 6
ION SUPPRESSON IN ESI AND APCI
Courtesy of LGC.com

## Slide 7
Background adduct ions of Esi over APCI at “the same LC-MS conditions”
ESi
apci

## Slide 8
Background adduct ions of Esi over APCI at “the same condition”
ESi
apci

## Slide 9
Chromatograms of Esi and APCI at “the same LC-MS conditions”
ESi
apci

## Slide 10
LC-MS chromatograms and MS spectra of Esi and APCI at “the same LC-MS conditions”
ESi
apci
apci
ESi
W.J. Cao, 2014
W.J. Cao, 2014

## Slide 11
APCI-MS scan spectrum of DNBP
W.J. Cao, 2014
W.J. Cao, 2014

## Slide 12
APCI Product ions of M+H+
Ready for MRM
W.J. Cao, 2014
W.J. Cao, 2014

## Slide 13
ESI-MS Scan spectrum of DNBP at the “same condition”
May not be able to do fast FIA or Infusion optimization
W.J. Cao, 2014

## Slide 14
ESI PRODUCT ION SPECTRUM OF M+Na+
W.J. Cao, 2014

## Slide 15
AGILENT JET-STREAM ESI
Courtesy of Agilent.com

## Slide 16
AJS-ESI SPECTRUM OF DNBP
W.J. Cao, 2014
W.J. Cao, 2014
Figure I. ESI produce many metal adduct ions which can’t be used for MRM analysis.

## Slide 17
AJS-ESI PRODUCT IONs OF 2M+Na+ At LOW CE
W.J. Cao, 2014

## Slide 18
AJS-ESI PRODUCT IONS OF 2M+Na+ At HIGHER CE
W.J. Cao, 2014
W.J. Cao, 2014

## Slide 19
Chromatogram of LC-UV @ 270 nm
Chromatogram of LC-MSD (APCI +)
Strong UV adsorption indicates aromatic groups
No UV peaks, nice series
Characterization of Polymer Additives by LC-APCI-MSD

## Slide 20
Dm=44
Nonylphenol Ethoxylates (NPE)
Alkylphenol Ethoxylates (APE)
Characterization of Polymer Additives by LC-APCI-MS

## Slide 21
Dm=28
Dm=28
Dm=42
Dm=28
Dm=42
Sorbitan Di-Laurate
Sorbitan Mono-Laurate
Characterization of Polymer Additives by LC-APCI-MSD
Wenjie Cao, 2007

## Slide 22
Quantitation by UHPLC-APCI-MS/ms
Wenjie Cao, November 2011

## Slide 23
Wenjie Cao, November 2011
Quantitation by UHPLC-APCI-MS/MS

## Slide 24
EI, ESI, CI, APCI, A trouble shooting story
“Perfect” reasoning with publication backup
What is that 106 ion?

## Slide 25
Perfect reasoning, but fundamentally wrong A trouble shooting story Mislead by publications

## Slide 26
Perfect reasoning, but fundamentally wrong A trouble shooting story Mislead by publications

## Slide 27
EI, ESI, CI, APCI, One more publication on APCI
Use air in APCI?
APCI needs “greater C of additives in mobile phase?

## Slide 28
Summary of ESI and APCI for lc-ms
Summary of Pros and Cons of ESI and APCI |  |  | 
ESI |  | APCI | 
Pros | Cons | Pros | Cons *
Multiple charged ions for high MW chemicals | Could be difficult in MS interpretation | Proton exchange, easier to get M+1, M-1 | Can't handle chemicals which can’t do proton-exchange.
Metal ions for confirmation of M? | Metal adducts | No multiple-charged ions, easier for identification | MW range is usually less than 1500
 | May mess up with MRM | No metal ions, easier for identification | 
 | Ion suppression could be more serious | Less ion suppression | 
 | More background ions, noisy baseline | No metal ions to mess up MRM | 
Ionic chemicals | More matrix dependent | Neater background | 
Popular, more publications and references |  |  | Less good publications and references.
Low T, good for thermally unstable chemicals | More adducts and clusters | High T, Less cluster/adduct ions | Thermally unstable chemicals
 | Can't handle high flow rate | Can ionize less polar chemicals | 
 | Ion suppression could be more serious | High T, can handle higher LC-flow rate | 
 |  | * may not be APCI issues | 

## Slide 29
Thank you!
Questions?