## Slide 1: THE EFFECT OF ANTICOAGULANT TYPE ON LEVOFLOXACIN ANALYSIS IN PLASMA USING HIGH PERFORMANCE LIQUID CHROMATOGRAPHY – PHOTODIODE ARRAY
Yahdiana Harahap, Agus Imam Bahaudin, Harmita
Bioavailability/Bioequivalence Laboratory
Faculty of Pharmacy, Universitas Indonesia
  Analytical & Bioanalytical Techniques
Valencia, Spain 2015

## Slide 2: INTRODUCTION
Levofloxacin is a broad spectrum antibiotic and is one of the compulsory drugs for bioequivalence testing, so that determination of its level in plasma is needed to be done for monitoring and ensuring its quality
Plasma is the most used biological matrix in bioanalysis and when it is used in an analysis, one of the important factors to be noticed is the types of anticoagulant being used.
Anticoagulant is used to avoid coagulation in handling the blood sample.
Ethylene Diamine Tetraacetate (EDTA), heparin, and citrate are anticoagulants often used in drug analysis in human blood

## Slide 3: BACKGROUND
Analyte Molecule
Metabolism Profile
Clinical Parameter
EDTA (Ethylene Diamine Tetra Acetate)
Chelating agent Ca2+
Heparin
Anti-Trombin III
Citrate / Citric Phosphate Dextrose (CPD)
Chelating agent Ca2+
Type of Anticoagulant

## Slide 4: THE OBJECTIVE OF RESEARCH
To develop and validate levofloxacin method analysis in human plasma by HPLC-Photodiode array and observing the effect of anticoagulant types to parameters stability, recovery, and the chromatogram profile .
The validated method will be applied to pilot bioequivalence study with six healthy subjects which were administered two 500 mg levofloxacin formulations.

## Slide 6
Instruments and Chromatography condition
Preparation of Standard Calibrators
Sample Preparation
Method Validation
Comparison of levofloxacin analysis using 3 types of anticoagulant
Application to pilot bioequicalence study
Methodology

## Slide 7: Instruments
The HPLC (Waters Alliance 2695) consisted of degasser, HPLC pump, data processor, oven column, and photo diode array detector (Waters 2996).

## Slide 8
Column : C18 SunfireTM (5µm, 250 x 4,6 mm)
Column temperature : 45°C
Mobile Phase : 0.5% triethylamine pH 3.0 - acetonitrile (88:12)
Flow rate : 1.25 mL/minute
Detection : 293nm
Chromatographic Condition

## Slide 9: Preparation of Standard
The standard curve in the concentration range of 50-10000 ng/mL was used in this study. The calibration samples were prepared at nine different concentration levels with blank and zero.

## Slide 10: Blood Collection and Plasma Isolation
Six Healthy Subjects
Heparin Plasma
EDTA Plasma
Citrate Plasma
Vacutainer Li-Heparin
Vacutainer K3EDTA
centrifugation
centrifugation

## Slide 11: Sample Preparation
300 µL Plasma with Levofloxacin
(+) 50 µL Ciprofloxacin HCl 100.0 µg/mL
Add 450 µL Methanol
Shake with Vortex
for 3 mins
Centrifugation 10,000 g
for 10 mins
20 µL aliquot injected into HPLC
vortex

## Slide 12: Method Validation of Levofloxacin in Plasma
Heparin Plasma
EDTA Plasma
Citrate Plasma
Partial Validation:
Accuracy
Precision
-Stability Test
Full Validation:
LLOQ
Calibration Curve/Linearity
Selectivity
Carry-over
Accuracy & Precision
Recovery
Dilution integrity
Stability Test

## Slide 13: LLOQ
Conc. ng/mL | Area |  | PAR | Measured Concentration |  |  |  | 
 | Analyte | IS |  | ng/mL | Average | SD | % CV | % diff
50.00 | 2018 | 221269 | 0.0091 | 55.18 | 50.63 | 5.35 | 10.57 | 10.37
 | 1797 | 222022 | 0.0081 | 48.34 |  |  |  | -3.32
 | 1795 | 202890 | 0.0088 | 53.36 |  |  |  | 6,73
 | 1446 | 201289 | 0.0072 | 42.28 |  |  |  | -15.45
 | 1811 | 202550 | 0.0089 | 53.99 |  |  |  | 7.98
25.00 | 626 | 213044 | 0.0029 | 13.98 | 15.53 | 3.41 | 21.94 | -44.09
 | 857 | 230882 | 0.0037 | 19.13 |  |  |  | -23.46
 | 877 | 234988 | 0.0037 | 19.27 |  |  |  | -22,92
 | 578 | 206921 | 0.0028 | 13.01 |  |  |  | -47.95
 | 556 | 207467 | 0.0027 | 12.26 |  |  |  | -50.98

## Slide 14: Calibration Curve
Concentration (ng/mL) | Area |  | PAR | Measured Concentration (ng/mL) | % diff
 | Analyte | IS |  |  | 
0.00 |  | 207181 | 0.0000 | 0.00 | 
50.00 | 1525 | 207732 | 0.0073 | 43.33 | -13.35
100.00 | 3704 | 219318 | 0.0169 | 106.97 | 6.97
200.00 | 6474 | 225071 | 0.0288 | 186.12 | -6.94
500.00 | 15311 | 199824 | 0.0766 | 505.13 | 1.03
1000.00 | 31286 | 199670 | 0.1567 | 1038.81 | 3.88
5000.00 | 156726 | 210453 | 0.7447 | 4958.30 | -0.83
10000.00 | 315374 | 209742 | 1.5036 | 10016.95 | 0.17
Slope (b) | Intercept (a) | R | R2
0.0002 | 0.0008 | 1.0000 | 1.0000

## Slide 15: Calibration Curve Inter-day
Conc. (ng/mL) | 0 | 50 | 100 | 200 | 500 | 1000 |  | 5000 |  |  | 10000 | R | Slope (b) | Intercept (a)
No. | Measured Concentration (ng/ml) |  |  |  |  |  |  |  |  |  |  |  |  | 
1 | 0.00 | 59.63 | 106.30 | 225.66 | 440.48 |  | 1025.36 |  | 4998.80 | 10000.42 |  | 1.0000 | 0.0001 | 0.0007
2 | 0.00 | 54.73 | 113.21 | 199.70 | 512.97 |  | 1042.01 |  | 4907.66 | 10041.17 |  | 0.9999 | 0.0001 | 0.0032
3 | 0.00 | 55.09 | 112.16 | 182.05 | 474.47 |  | 1008.02 |  | 4998.77 | 10001.30 |  | 1.0000 | 0.0002 | -0.0027
4 | 0.00 | 53.32 | 104.59 | 211.51 | 449.55 |  | 1040.93 |  | 5033.07 | 9981.60 |  | 1.0000 | 0.0001 | 0.0027
5 | 0.00 | 42.89 | 90.47 | 170.60 | 468.31 |  | 911.97 |  | 5294.21 | 9864.00 |  | 0.9994 | 0.0002 | -0.0011
Average | 0.00 | 56.48 | 110.56 | 202.47 | 475.98 |  | 1025.13 |  | 4968.41 | 10014.30 |  | 0.9998 | 0.0001 | 0.0006
SD | 0.00 | 2.73 | 3.72 | 21.94 | 36.27 |  | 16.99 |  | 52.61 | 23.28 |  |  |  | 
CV (%) | 0.00 | 4.84 | 3.37 | 10.83 | 7.62 |  | 1.66 |  | 1.06 | 0.23 |  |  |  | 

## Slide 16: Selectivity
Conc. (ng/mL) |  | Area |  | PAR | Measured Concentration |  |  |  | % diff
 |  | Analyte | IS |  | ng/mL | Average | SD | % CV | 
50.00 | A | 2018 | 221269 | 0.0091 | 58.41 | 53.05 | 5.07 | 9.56 | 16.81
 |  | 1675 | 200094 | 0.0084 | 51.59 |  |  |  | 3.18
 | B | 1797 | 222022 | 0.0081 | 49.07 |  |  |  | -1.86
 |  | 1814 | 197173 | 0.0092 | 59.13 |  |  |  | 18.27
 | C | 1795 | 202890 | 0.0088 | 55.92 |  |  |  | 11.85
 |  | 1716 | 196319 | 0.0087 | 54.96 |  |  |  | 9.91
 | D | 1446 | 201289 | 0.0072 | 40.79 |  |  |  | -18.42
 |  | 1542 | 188658 | 0.0082 | 49.79 |  |  |  | -0.41
 | E | 1811 | 202550 | 0.0089 | 56.78 |  |  |  | 13.55
 |  | 1684 | 201939 | 0.0083 | 51.30 |  |  |  | 2.60
 | F | 1893 | 213495 | 0.0089 | 56.10 |  |  |  | 12.20
 |  | 1996 | 234873 | 0.0085 | 52.75 |  |  |  | 5.50

## Slide 17: Overlay Plasma LLOQ-Blank
Blank Plasma
Plasma LLOQ

## Slide 18: Accuracy, Precision, and Recovery
Sample
5 repl
LLOQ
QCL = 3 x LLOQ
QCM = ± 50% Calibration Curve
QCH = 75% ULOQ
Within-run (intra-day)
Parameter:
Accuracy  % diff
Precision  % CV
Rec  Absolute Recovery and Modification
Between-run (intra-day)
Citrate Plasma
Heparin
Plasma
EDTA
Plasma
Min. 2 different days

## Slide 19: Result of Accuracy and Precision of Citrate Plasma
Conc. ng/mL | Day | Measured Conc. | % CV | % diff | % CV
50.00 | 1 | 53.29 | 10.85 | 6.59 | 8.94
 |  | 46.02 |  | -7.96 | 
 |  | 44.65 |  | -10.71 | 
 |  | 40.59 |  | -18.83 | 
 |  | 51.17 |  | 2.35 | 
 | 2 | 52.18 | 17.16 | 4.35 | 
 |  | 40.20 |  | -19.61 | 
 |  | 40.92 |  | -18.16 | 
 |  | 56.48 |  | 12.96 | 
 |  | 58.06 |  | 16.12 | 
 | 3 | 56.10 | 4.67 | 12.21 | 
 |  | 55.23 |  | 10.47 | 
 |  | 59.36 |  | 18.72 | 
 |  | 52.17 |  | 4.34 | 
 |  | 56.87 |  | 13.75 | 
Conc. ng/mL | Day | Measured Conc. | % CV | % diff | % CV
150.00 | 1 | 166.70 | 3.23 | 11.14 | 7.16
 |  | 155.73 |  | 3.82 | 
 |  | 162.54 |  | 8.36 | 
 |  | 169.85 |  | 13.23 | 
 |  | 164.19 |  | 9.46 | 
 | 2 | 128.23 | 8.81 | -14.52 | 
 |  | 150.25 |  | 0.16 | 
 |  | 132.66 |  | -11.56 | 
 |  | 156.56 |  | 4.37 | 
 |  | 152.30 |  | 1.53 | 
 | 3 | 141.86 | 4.61 | -5.43 | 
 |  | 147.53 |  | -1.64 | 
 |  | 139.79 |  | -6.81 | 
 |  | 145.03 |  | -3.31 | 
 |  | 157.11 |  | 4.74 | 

## Slide 20: Result of Accuracy and Precision of Citrate Plasma
Conc. ng/mL | Day | Measured Conc. | % CV | % diff | % CV
5000 | 1 | 4294.73 | 7.69 | -14.11 | 3.72
 |  | 4949.63 |  | -1.01 | 
 |  | 5174.05 |  | 3.48 | 
 |  | 4500.77 |  | -9.98 | 
 |  | 4529.02 |  | -9.42 | 
 | 2 | 4336.18 | 3.09 | -13.28 | 
 |  | 4408.63 |  | -11.83 | 
 |  | 4504.54 |  | -9.91 | 
 |  | 4635.21 |  | -7.30 | 
 |  | 4294.67 |  | -14.11 | 
 | 3 | 4475.16 | 3.30 | -10.50 | 
 |  | 4254.95 |  | -14.90 | 
 |  | 4265.83 |  | -14.68 | 
 |  | 4294.45 |  | -14.11 | 
 |  | 4576.37 |  | -8.47 | 
Conc. ng/mL | Day | Measured Conc. | % CV | % diff | % CV
7500 | 1 | 8579.07 | 0.39 | 14.39 | 12.11
 |  | 8563.39 |  | 14.18 | 
 |  | 8551.07 |  | 14.01 | 
 |  | 8521.93 |  | 13.63 | 
 |  | 8496.01 |  | 13.28 | 
 | 2 | 6847.25 | 2.23 | -8.70 | 
 |  | 6833.94 |  | -8.88 | 
 |  | 7011.39 |  | -6.51 | 
 |  | 6591.29 |  | -12.12 | 
 |  | 6766.47 |  | -9.78 | 
 | 3 | 6922.07 | 3.53 | -7.71 | 
 |  | 7061.70 |  | -5.84 | 
 |  | 7570.93 |  | 0.95 | 
 |  | 7309.02 |  | -2.55 | 
 |  | 7091.72 |  | -5.44 | 

## Slide 21: Result of Accuracy and Precision of Heparin Plasma
Conc. ng/mL | Day | Measured Conc. | % CV | % diff | % CV
50.00 | 1 | 47.51 | 5.99 | -4.97 | 2.27
 |  | 47.71 |  | -4.58 | 
 |  | 53.57 |  | 7.14 | 
 |  | 52.57 |  | 5.14 | 
 |  | 53.28 |  | 6.55 | 
 | 2 | 50.48 | 5.01 | 0.97 | 
 |  | 50.50 |  | 1.00 | 
 |  | 53.96 |  | 7.93 | 
 |  | 56.02 |  | 12.03 | 
 |  | 55.51 |  | 11.03 | 
 | 3 | 52.83 | 7.28 | 5.66 | 
 |  | 52.99 |  | 5.98 | 
 |  | 55.15 |  | 10.29 | 
 |  | 54.35 |  | 8.70 | 
 |  | 45.61 |  | -8.77 | 
Conc. ng/mL | Day | Measured Conc. | % CV | % diff | % CV
150.00 | 1 | 132.01 | 10.91 | -11.99 | 6.13
 |  | 129.84 |  | -13.44 | 
 |  | 137.67 |  | -8.22 | 
 |  | 150.61 |  | 0.41 | 
 |  | 167.52 |  | 11.68 | 
 | 2 | 152.31 | 2.70 | 1.54 | 
 |  | 146.10 |  | -2.60 | 
 |  | 156.36 |  | 4.24 | 
 |  | 155.99 |  | 3.99 | 
 |  | 153.38 |  | 2.26 | 
 | 3 | 133.35 | 3.31 | -11.10 | 
 |  | 139.08 |  | -7.28 | 
 |  | 140.86 |  | -6.09 | 
 |  | 131.06 |  | -12.63 | 
 |  | 131.70 |  | -12.20 | 

## Slide 22: Result of Accuracy and Precision of Heparin Plasma
Conc. ng/mL | Day | Measured Conc. | % CV | % diff | % CV
5000 | 1 | 4348.19 | 0.88 | -13.04 | 9.19
 |  | 4333.78 |  | -13.32 | 
 |  | 4362.10 |  | -12.76 | 
 |  | 4434.01 |  | -11.32 | 
 |  | 4363.40 |  | -12.73 | 
 | 2 | 4464.63 | 1.91 | -10.71 | 
 |  | 4276.02 |  | -14.48 | 
 |  | 4264.94 |  | -14.70 | 
 |  | 4282.25 |  | -14.35 | 
 |  | 4318.77 |  | -13.62 | 
 | 3 | 4916.75 | 2.98 | -1.67 | 
 |  | 5042.10 |  | 0.84 | 
 |  | 4954.71 |  | -0.91 | 
 |  | 5181.81 |  | 3.64 | 
 |  | 5273.00 |  | 5.46 | 
Conc. ng/mL | Day | Measured Conc. | % CV | % diff | % CV
7500 | 1 | 7619.34 | 5.40 | 1.59 | 5.32
 |  | 7936.15 |  | 5.82 | 
 |  | 7544.60 |  | 0.59 | 
 |  | 6833.72 |  | -8.88 | 
 |  | 7573.50 |  | 0.98 | 
 | 2 | 8357.92 | 3.98 | 11.44 | 
 |  | 8509.14 |  | 13.46 | 
 |  | 7667.75 |  | 2.24 | 
 |  | 8349.55 |  | 11.33 | 
 |  | 8309.99 |  | 10.80 | 
 | 3 | 8293.64 | 3.49 | 10.58 | 
 |  | 8373.96 |  | 11.65 | 
 |  | 7739.40 |  | 3.19 | 
 |  | 8301.36 |  | 10.68 | 
 |  | 8473.62 |  | 12.98 | 

## Slide 23: Result of Accuracy and Precision of EDTA Plasma
Conc. ng/mL | Day | Measured Conc. | % CV | % diff | % CV
50.00 | 1 | 53.06 | 4.11 | 6.12 | 2.70
 |  | 51.63 |  | 3.25 | 
 |  | 55.71 |  | 11.42 | 
 |  | 56.73 |  | 13.47 | 
 |  | 56.40 |  | 12.79 | 
 | 2 | 53.04 | 5.20 | 6.07 | 
 |  | 53.44 |  | 6.87 | 
 |  | 49.49 |  | -1.02 | 
 |  | 51.71 |  | 3.43 | 
 |  | 57.02 |  | 14.04 | 
 | 3 | 48.73 | 5.74 | -2.54 | 
 |  | 52.92 |  | 5.83 | 
 |  | 49.91 |  | -0.18 | 
 |  | 56.39 |  | 12.79 | 
 |  | 51.37 |  | 2.74 | 
Conc. ng/mL | Day | Measured Conc. | % CV | % diff | % CV
150.00 | 1 | 132.82 | 7.43 | -11.45 | 0.97
 |  | 145.76 |  | -2.82 | 
 |  | 160.28 |  | 6.85 | 
 |  | 158.73 |  | 5.82 | 
 |  | 149.74 |  | -0.17 | 
 | 2 | 138.60 | 7.95 | -7.60 | 
 |  | 156.38 |  | 4.25 | 
 |  | 160.91 |  | 7.27 | 
 |  | 133.13 |  | -11.25 | 
 |  | 145.65 |  | -2.90 | 
 | 3 | 141.55 | 6.83 | -5.63 | 
 |  | 133.80 |  | -10.80 | 
 |  | 148.83 |  | -0.78 | 
 |  | 160.56 |  | 7.04 | 
 |  | 150.44 |  | 0.29 | 

## Slide 24: Result of Accuracy and Precision of EDTA Plasma
Conc. ng/mL | Day | Measured Conc. | % CV | % diff | % CV
5000 | 1 | 4934.47 | 5.78 | -1.31 | 6.47
 |  | 5145.11 |  | 2.90 | 
 |  | 4439.22 |  | -11.22 | 
 |  | 4823.51 |  | -3.53 | 
 |  | 5100.48 |  | 2.01 | 
 | 2 | 4757.35 | 4.72 | -4.85 | 
 |  | 4401.42 |  | -11.97 | 
 |  | 4287.12 |  | -14.26 | 
 |  | 4261.61 |  | -14.77 | 
 |  | 4285.99 |  | -14.28 | 
 | 3 | 4901.06 | 1.75 | -1.98 | 
 |  | 4941.48 |  | -1.17 | 
 |  | 4870.09 |  | -2.60 | 
 |  | 5062.68 |  | 1.25 | 
 |  | 5049.33 |  | 0.99 | 
Conc. ng/mL | Day | Measured Conc. | % CV | % diff | % CV
7500 | 1 | 8250.49 | 1.35 | 10.01 | 6.54
 |  | 8380.35 |  | 11.74 | 
 |  | 8414.35 |  | 12.19 | 
 |  | 8546.61 |  | 13.95 | 
 |  | 8494.73 |  | 13.26 | 
 | 2 | 7246.72 | 7.50 | -3.38 | 
 |  | 7010.43 |  | -6.53 | 
 |  | 8446.75 |  | 12.62 | 
 |  | 7322.54 |  | -2.37 | 
 |  | 7770.12 |  | 3.60 | 
 | 3 | 7414.01 | 4.71 | -1.15 | 
 |  | 7074.69 |  | -5.67 | 
 |  | 7601.88 |  | 1.36 | 
 |  | 8039.14 |  | 7.19 | 
 |  | 7392.79 |  | -1.43 | 

## Slide 25: Recovery
Concentration | Citrate |  | Heparin |  | EDTA | 
 | % Rec | %CV | % Rec | %CV | % Rec | %CV
QCL | 80.35 | 6.29 | 84.20 | 2.09 | 75.57 | 9.34
 | 90.86 |  | 82.37 |  | 86.22 | 
 | 83.85 |  | 80.75 |  | 90.94 | 
QCM | 88.22 | 3.66 | 77.87 | 6.39 | 69.44 | 8.09
 | 83.20 |  | 86.77 |  | 80.32 | 
 | 89.09 |  | 87.55 |  | 80.03 | 
QCH | 103.75 | 3.82 | 100.07 | 7.10 | 90.26 | 8.39
 | 100.55 |  | 101.10 |  | 103.13 | 
 | 108.47 |  | 113.45 |  | 105.98 | 
IS | 43.40 | 2.01 | 46.10 | 9.56 | 50.49 | 5.80
 | 42.86 |  | 48.52 |  | 47.66 | 
 | 41.72 |  | 55.32 |  | 44.96 | 

## Slide 26: Dilution Integrity
Sample
Concentration above ULOQ
Half Dilution
Quarter Dilution
Objective:
To check whether the dilution during in-study analysis is reliable or not especially when samples are over ULOQ.

## Slide 27: Dilution Integrity
Half Dilution |  |  |  |  | 
Conc. | Day | Meas Conc | % CV | % diff | % CV
7500 | 1 | 7405.31 | 2.73 | -1.26 | 6.51
 |  | 7868.28 |  | 4.91 | 
 |  | 7789.45 |  | 3.86 | 
 |  | 7874.63 |  | 5.00 | 
 |  | 7932.50 |  | 5.77 | 
 | 2 | 7255.18 | 3.06 | -3.26 | 
 |  | 7072.10 |  | -5.71 | 
 |  | 7178.69 |  | -4.28 | 
 |  | 6954.84 |  | -7.27 | 
 |  | 6705.20 |  | -10,60 | 
 | 3 | 6885.55 | 1.61 | -8.19 | 
 |  | 7055.88 |  | -5.92 | 
 |  | 6882,88 |  | -8.23 | 
 |  | 6745.84 |  | -10.06 | 
 |  | 6923.86 |  | -7.68 | 
Quarter Dilution |  |  |  |  | 
Conc. | Day | Meas Conc | % CV | % diff | % CV
3750 | 1 | 3486.88 | 3.11 | -7.02 | 4.36
 |  | 3611.49 |  | -3.69 | 
 |  | 3647.82 |  | -2.72 | 
 |  | 3405.76 |  | -9.18 | 
 |  | 3424.06 |  | -8.69 | 
 | 2 | 3222.80 | 1.01 | -14.06 | 
 |  | 3245.27 |  | -13.46 | 
 |  | 3299.70 |  | -12.01 | 
 |  | 3218.47 |  | -14.17 | 
 |  | 3258.33 |  | -13.11 | 
 | 3 | 3220.26 | 2.78 | -14.13 | 
 |  | 3426.82 |  | -8.62 | 
 |  | 3194.40 |  | -14.82 | 
 |  | 3254.09 |  | -13.22 | 
 |  | 3295.14 |  | -12.13 | 
  Analytical & Bioanalytical Techniques

## Slide 28: Carry-Over
No. | Area on blank plasma |  | Area on LLOQ Plasma Sample |  | % Carry-over | 
 | Analyte | IS | Analyte | IS | Analyte | IS
1 | 43 | 92 | 2018 | 221269 | 2.13% | 0.04%
2 | 40 | 108 | 1797 | 222022 | 2.23% | 0.05%
3 | 35 | 79 | 1795 | 202890 | 1.95% | 0.04%
4 | 71 | 111 | 1446 | 201289 | 4.91% | 0.06%
5 | 41 | 90 | 1811 | 202550 | 2.26% | 0.04%
Average |  |  |  |  | 2.70% | 0.05%
Requirement % carry-over |  |  |  |  | < 20% | < 5 %

## Slide 29: Stability Test
Freeze-Thaw Stability
Long Term Stability
Stock Solution Stability
Short Term Stability
Autosampler Stability
Citrate
Plasma
Heparin
Plasma
EDTA
Plasma

## Slide 30: Freeze and Thaw Stability
Citrate Plasma |  | 
Conc. | Cycle | % diff
QCL | 0 Cycle | 11.98
 |  | -3.88
 |  | -7.82
 | 3 Cycle | 14.68
 |  | 12.42
 |  | -2.64
QCH | 0 Cycle | 9.93
 |  | 13.98
 |  | 12.81
 | 3 Cycle | 13.79
 |  | 14.16
 |  | 11.07
Heparin Plasma |  | 
Conc. | Cycle | % diff
QCL | 0 Cycle | -6.20
 |  | -2.93
 |  | -3.66
 | 3 Cycle | 8.45
 |  | 11.61
 |  | 5.57
QCH | 0 Cycle | 14.74
 |  | 14.39
 |  | 14.53
 | 3 Cycle | 14.88
 |  | 14.71
 |  | 13.93
EDTA Plasma |  | 
Conc. | Cycle | % diff
QCL | 0 Cycle | 3.90
 |  | 3.82
 |  | -2.14
 | 3 Cycle | -8.49
 |  | -5.11
 |  | -10.68
QCH | 0 Cycle | -14.92
 |  | 14.57
 |  | 13.39
 | 3 Cycle | 14.03
 |  | 14.57
 |  | 14.25

## Slide 31: Short Term Stability
Time | % diff Citrate |  | % diff Heparin |  | % diff EDTA | 
 | QCL | QCH | QCL | QCH | QCL | QCH
Hour -0 | 11.98 | 9.93 | -6.20 | 14.74 | 3.90 | -14.92
 | -3.88 | 13.98 | -2.93 | 14.39 | 3.82 | 14.57
 | -7.82 | 12.81 | -3.66 | 14.53 | -2.14 | 13.39
Hour-7 | 2.01 | 5.22 | 6.06 | 12.64 | -4.41 | 5.44
 | 13.60 | 9.00 | -1.80 | 14.83 | -13.56 | 9.83
 | -11.70 | -2.81 | -12.63 | 14.00 | 11.88 | 14.21
Hour-24 | -12.80 | -11.03 | 11.74 | 11.22 | 2.06 | 14.65
 | -13.99 | 8.66 | 6.82 | 5.11 | 0.83 | 14.69
 | -1.54 | 7.40 | 8.16 | -1.43 | -5.39 | 14.32

## Slide 32: Long Term Stability
Time | % diff Citrate |  | % diff Heparin |  | % diff EDTA | 
 | QCL | QCH | QCL | QCH | QCL | QCH
Day-1 | 11.98 | 9.93 | -6.20 | 14.74 | 3.90 | -14.92
 | -3.88 | 13.98 | -2.93 | 14.39 | 3.82 | 14.57
 | -7.82 | 12.81 | -3.66 | 14.53 | -2.14 | 13.39
Day-7 | -0.11 | -10.37 | 13.24 | -3.91 | 0.37 | -14.74
 | 9.45 | -12.25 | 14.15 | -1.78 | -3.22 | -7.14
 | 4.47 | -13.64 | 14.11 | -7.59 | -1.75 | -10.40
Day-14 | -6.07 | -14.18 | -5.09 | -12.40 | -12.72 | -9.80
 | 1.53 | -13.41 | -3.08 | -11.63 | -14.18 | -7.65
 | -1.31 | -14.68 | -7.95 | -5.95 | -14.09 | -12.14
Day-23 | 4.89 | -13.38 | 8.84 | -7.17 | -11.80 | -3.67
 | 5.06 | -11.94 | -0.67 | -8.38 | 1.62 | -14.35
 | 10.64 | -14.88 | -4.74 | -8.21 | -6.31 | -6.00
Day-31 | 13.04 | -0.53 | -9.14 | -10.96 | -12.04 | -5.84
 | -1.48 | -2.25 | 2.66 | 3.03 | -13.11 | 5.42
 | 10.39 | 0.85 | -7.11 | 11.55 | -13.85 | -3.43

## Slide 33: Autosampler Stability
Time (hour) | % diff Citrate |  | % diff Heparin |  | % diff EDTA | 
 | QCL | QCH | QCL | QCH | QCL | QCH
0 | 11.98 | 9.93 | -6.20 | 14.74 | 3.90 | -14.92
 | -3.88 | 13.98 | -2.93 | 14.39 | 3.82 | 14.57
 | -7.82 | 12.81 | -3.66 | 14.53 | -2.14 | 13.39
24 | 10.16 | 6.06 | 5.44 | 14.41 | -1.78 | -13.19
 | 4.42 | 10.70 | -3.85 | 14.83 | -12.98 | 12.74
 | 9.07 | 12.64 | -10.96 | 10.99 | 14.22 | 12.25

## Slide 34: Comparison of Levofloxacin Analysis Using 3 Types of Anticoagulant

## Slide 35
Heparin Plasma Blank
EDTA Plasma Blank
Citrate Plasma Blank

## Slide 36
Parameter Analysis |  | Plasma Type |  |  |  |  |  | Difference | PValue
 |  | Citrate |  | Heparin |  | EDTA |  |  | 
Recovery (% ± SD) |  |  |  |  |  |  |  |  | 
1.QCL | 85.02 ± 5.35 |  | 82.44 ± 1.72 |  | 84.24 ± 7.87 |  | There is no significant difference |  | 0.849
2.QCH | 104.25 ± 3.99 |  | 104.87 ± 74.88 |  | 99.79 ± 8.38 |  |  |  | 
 |  |  |  |  |  |  |  |  | 0.635
Peak Area Ratio (± SD) |  |  |  |  |  |  |  |  | 
1.QCL | 0.0211 ± 0.0030 |  | 0.0221 ± 0.0024 |  | 0.0177 ± 0.0020 |  | Citrate-EDTA; Heparin-EDTA |  | 0.000*
2.QCM | 0.6174 ± 0.0714 |  | 0.6896 ± 0.0550 |  | 0.6896 ± 0.0564 |  | Sitrat-Heparin; Sitrat-EDTA |  | 0.020*
3.QCH | 1.0218 ± 0.0569 |  | 1.2012 ± 0.0723 |  | 1.1396 ± 0.1092 |  | Sitrat-Heparin; Sitrat-EDTA |  | 0.000*
P > 0,05 ; no significant difference
P < 0,05 ; significant different
* Kruskal Wallis and Mann-Whitney Test

## Slide 37
Parameter Analysis | Type of Plasma |  | 
 | Citrate | Heparin | EDTA
Stability at temperature of ±25°C (short term) |  |  | 
QCL | 24 h | 24 h | 24 h
QCH | 24 h | 24 h | 24 h
Stability post preparation (in autosampler) |  |  | 
QCL | 24 h | 24 h | 24 h
QCH | 24 h | 24 h | 24 h
Stability freeze-thaw cycle |  |  | 
QCL | 3 cycle | 3 cycle | 3 cycle
QCH | 3 cycle | 3 cycle | 3 cycle
Stability at temperature of -20°C (Long Term) |  |  | 
QCL | 31 days | 31 days | 31 days
QCH | 31 days | 31 days | 31 days

## Slide 38: Application to Pilot Bioequivalence Study

## Slide 39: Pilot Bioequivalence study
The validated method was applied to the pilot bioequivalence study of levofloxacin after a single oral administration of 500 mg levofloxacin film-coated tablet to 6 healthy Indonesian subjects.

## Slide 40
Study
Protocol
Subject screening
6 healthy subjects
Blood sampling pre dose, 0.25, 0.5, 1, 1.5, 2, 2.5, 3, 4, 6, 8, 12, 24 hours (one week wash-out period)
9 mL of Blood
Plasma was frozen at -20C
Plasma isolation using CPD-A
Committee of The Medical Research Ethics of the Faculty of Medicine, Universitas Indonesia
Signed
informed
consent
Validated method
Linear 50-10000 ng/mL
Blood pressure, heart
rate, body temperature and adverse events were monitored during blood sampling.
passed
single-dose, open-label,
randomized,
two-way crossover
fasting
Notes: 40

## Slide 41: Pharmacokinetic evaluation
Mean plasma concentration-time profiles of levofloxacin after a single dose of two 500 mg levofloxacin film-coated tablets of two different formulations

## Slide 42
Chromatogram of QCL
Chromatogram of QCM
Chromatogram of QCH
Chromatogram of Sample

## Slide 43: CONCLUSION
  Bioanalysis method has fulfilled the validation criteria, which was linear at 50 – 10000 ng/ml
  There was no significant difference in stability and recovery of levofloxacin in plasma (p > 0.05), however there was significant difference (p < 0.05) in PAR value.
  The method has been successfully applied for pilot bioequivalence study with six Indonesian healthy subjects using CPD-A as anticoagulant after administration of 500mg levofloxacin.

## Slide 44: ACKNOWLEDGEMENTS
This study was supported by the Directory of Research and Public Service Universitas Indonesia.

## Slide 45
Thank You
Universitas Indonesia