Identifying risks in a flexible fetoscope during the design phase by applying a modified Delphi method for failure mode and effect analysis

Phokee W1, Pintellon L2, Vander Poorten E3, Chatpun S1

1Department of Biomedical Sciences and Biomedical Engineering, Faculty of Medicine, Prince of Songkla University, Thailand

2Centre for Industrial Management, Department of Mechanical Engineering, Faculty of Engineering Technology, KU Leuven, Belgium

3Department of Mechanical Engineering, Faculty of Engineering Technology, Group T Leuven Campus, KU Leuven, Belgium

Abstract:

Statement of the Problem: Failure mode and effect analysis (FMEA) has been widely used in healthcare to prevent adverse medical events and improve patient safety. However, gathering high-quality data from a diverse team might be problematic when using a standard FMEA approach. Especially the existence of a group bias and the excessive time consumption may be problematic. This study proposes the use of a modified Delphi technique for FMEA to circumvent these problems.

Methods: The proposed method combining FMEA with a modified Delphi method was applied to identify the risks of the flexible fetoscope during the design phase. The data collection included eleven multidisciplinary experts, consisting of 6 engineers from KU Leuven and 5obstetrician surgeons from Thailand. The experts were asked open-ended questions via an online individual interview at a time that was found convenient to them. They were allowed to review and edit their opinion throughout, before a facilitator created a summary report.

Findings: This study had a 100% participation rate. Experts provided their opinions without the risk of being dominated by others, and they were not informed of other information due to the anonymity method. Both steps are critical components of the modified Delphi technique. Sixty-five potential failure modes were found. Twenty-nine (45%) potential failure modes affected patient safety (mother and fetus). The failure mode related to the tip of the instrument was marked the most often (46%), followed by the handle (27%) and the flexible segment (27%) of the scope.

Conclusion and Significant: The modified Delphi technique is discovered to be useful and time efficient. This approach allows each expert the convenience of selecting a suitable time for a meeting, without a series of multiple meetings. The modified Delphi method also reduces the risk of group bias by anonymity in individual interviews.

Biography:

Wipharat Phokee is a doctoral student at the Department of Biomedical Sciences and Biomedical Engineering, Faculty of Medicine, Prince of Songkla University, Thailand. Prior to that, she obtained two bachelor’s degrees: one in occupational health and safety from Mahidol University, Thailand in 2000, and another in law from Sukhothai Thammathirat Open University, Thailand, in 2013. In 2005, she obtained a master’s degree in safety engineering from Kasetsart University, Thailand. She is currently also an associate professor at Khon Kaen University, Thailand. Her research interests include risk management, accident prevention, FMEA, and the Delphi method.

About Research Topic: FMEA is used to examine healthcare systems.1,2 FMEA tries to minimize adverse medical occurrences by identifying a product's or system's failure modes.2,3 A multidisciplinary expert committee comprised of experts in the field with diverse expertise in this process.4,5 FMEA requires a sequence of team meetings as well as the seek for consensus among members.1,6 Meetings with a large number of participants might bring aspects of group bias.7

About Institution:

Department of Biomedical Sciences and Biomedical Engineering, Faculty of Medicine, Prince of Songkla University, is located in the south of Thailand. There are 2 units in the department: Biomedical Sciences and Biomedical Engineering. The research in the department covers cancer biology, infectious diseases, biomaterials, biosensors, medical data analytics, biomechanics, and medical devices. There are 18 full time research faculty and around 60 graduate students in the department.

References:

1. 	Liu HC, Zhang LJ, Ping YJ, Wang L. Failure mode and effects analysis for proactive healthcare risk evaluation: A systematic literature review. J Eval Clin Pract. 2020;26(4):1320-1337. doi:10.1111/jep.13317

2. 	Frewen H, Brown E, Jenkins M, O’Donovan A. Failure mode and effects analysis in a paperless radiotherapy department. J Med Imaging Radiat Oncol. 2018;62(5):707-715. doi:10.1111/1754-9485.12762

3. 	Xu AY, Bhatnagar J, Bednarz G, et al. Failure modes and effects analysis (FMEA) for Gamma Knife radiosurgery. J Appl Clin Med Phys. 2017;18(6):152-168. doi:10.1002/acm2.12205

4. 	Martinez‐Marquez D, Terhaer K, Scheinemann P, Mirnajafizadeh A, Carty CP, Stewart RA. Quality by Design for industry translation: Three‐dimensional risk assessment failure mode, effects, and criticality analysis for additively manufactured patient‐specific implants. Eng Reports. 2020;2:1-24. doi:10.1002/eng2.12113

5. 	Varndell W, Fry M, Elliott D. Applying real-time Delphi methods: development of a pain management survey in emergency nursing. BMC Nurs. 2021;20:149. doi:10.1186/s12912-021-00661-9

6. 	Wu Z, Liu W, Nie W. Literature review and prospect of the development and application of FMEA in manufacturing industry. Int J Adv Manuf Technol. 2021;112:1409-1436. doi:10.1007/s00170-020-06425-0

7. 	Liu HC, Li P, You JX, Chen YZ. A novel approach for FMEA: Combination of interval 2-tuple linguistic variables and gray relational analysis. Qual Reliab Eng Int. 2015;31(5):761-772. doi:10.1002/qre.1633