ABSTRACT

An Assessment to Determine Total Scatter Factors for Photon Beam Linear Accelerators

Barrington Brevitt1, Barnswell Dennis2, Pankaj Patel3, Mitko Voutchkov4

Affiliations:

1,4Department of Physics, The University of the West Indies, Kingston 7, Jamaica.

2National Cancer Treatment Centre, Kingston 3, Jamaica.

3Houston Methodist Hospital, Houston, Texas.

Objectives:  To explore a simple and effective way to determine the total scatter of a C- Series linear accelerator.

Methodology: Measurements for this study were acquired using a Varian C-Series linear accelerator, with a 6MV photon beam, a Blue water phantom, 2 IBA CC13 ion chambers and an IBA CCU electrometer. Measurements were acquired for field sizes ranging from 5cm x 5cm to 40cm x 40cm, increasing field size by an increment of 5cm. Three readings were collected for each field size and averaged. All readings were done at a depth of 10cm to reduce the probability of measuring electron contamination in the photon beam. Measurements for Sc calculation were acquired in air using a build-up cap with a 3cm diameter.

Results: The average radiation dose measured increased as the field sizes increased. The maximum dose recorded for Sc was 2.33cGy while that for Scp was7.96cGy. There was a non-linear direct relationship between radiation dose measured, Scp and Sc calculated and field size. The maximum standard deviation in charge readings for Sc measurement was 1.18% which was recorded for the 35cm x 35cm field. The minimum standard deviation was 0.70% obtained with the 20cm x 20cm field. The proportion of the phantom scatter contribution to the total scatter, decreased exponentially with increases in field size. The largest contribution was identified with the 10cm x 10cm field while the lowest was identified with the 40cm x 40cm field.

Conclusion: It can be concluded that this method was effective in assessing the total scatter factor and its derivatives for field sizes ranging from 5cm x 5cm to 40cm x 40cm.

Keywords: Collimator Scatter; Phantom Scatter; Beam Fluence; PDD; Dmax; PTV; GTV

References

1. 	Podgorsak EB. Radiation Oncology Physics: Handbook for Teachers and Students. IAEA. 2005.

2. 	Grant, G.N, Lalor, G.C, Thomas JE. Thermoluminescence Dosimetry in the Gy Range , by P . Spanne. West Indian Med J. 2003;52(2):118–23.

3. 	F.M.Khan. The Physics Of Radiation Theraphy. 2010.

4. 	Tahmasebi Birgani MJ, Chegeni N, Behrooz MA, Bagheri M, Danyaei A, Shamsi A. An Analytical Method to Calculate Phantom Scatter Factor for Photon Beam Accelerators. Electron physician. 2017;9(1):3523–8.

5.	Griffith RL, Levi MR, Cartano D. The Photoelectric Effect. 2014.

6. 	Bushong S. Radiologic Science for Technologists. Tenth Edit. Radiologic Science for Technologists, 10/e. Elsevier; 2005.

7. 	Systems VM. Admin and Physics. Varian Medical System; 2018. p. 1028.

Biography

I am a final year PhD candidate at the University of the West Indies (UWI), Mona Campus, Kingston Jamaica doing research in Applied Physics. I am also pursuing a Post Graduate Certificate in University Teaching and Learning. I possess a Diploma in Sports Therapy, BSc in Diagnostic Imaging and an MSc in Medical Physics. I am currently employed as a Senior Medical Physicist and Radiation safety Officer at the Kingston Public Hospital and adjunct assistant lecturer at UWI. I am actively involved in radiation therapy treatment planning, and perform diagnostic and therapeutic quality assurance for radiation generating equipment. I conduct staff training and sensitize patients about their procedures. I have been involved in the medical use of radiation for the past 13 years; I have a passion for radiation safety and management and have published 6 research papers about same.