A regional level survey on acceptance test assessment in radio diagnostic instruments

C. Senthamil Selvan1 & 2 and C. S. Sureka2

1 & 2 Department of Medical Physics, Shri Vaishnav Vidyapeeth Vishwavidyalaya, Indore, Madhya Pradesh, India.

2 Department of Medical Physics, Bharathiar University, Coimbatore, Tamil Nadu, India.

Keywords: Diagnostic Radiology, Quality Assurance Tests, Dosimetry, Radiation Survey and Protection.

I. Abstract Body

The present study emphasizes the current status of quality assurance (QA) assessment on various radio-diagnostic instruments in and around southern parts of India. A pilot survey was conducted at the 230 instruments from 151 hospitals over a period of three years (2014–2017). This study includes as per Atomic Energy Regulatory Board, India (AERB) guidelines 16 mandatory tests were performed in X-ray tube such as congruence of radiation and optical field, central beam alignment, determination of focal spot size, linearity of exposure timer, accuracy of X-ray tube operating potential (kVp), linearity of X-ray tube current (mA), half value layer (HVL), total filtration, output consistency, measurement of computed tomography dose index (CTDI), slice thickness in CT, low and high contrast resolution, performance check of the image intensifier tube (in C-Arm and O-Arm), calibration of compression device (in Mammography), X-ray tube housing leakage radiation, and radiation survey throughout diagnostic department. In this work, we used calibrated PTW-NOMEX multimeter, pressurized ion chamber based survey meter, and other QA tools. Out of 230 instruments, 78% of the equipments were found good condition and the remaining 22 % of the equipments with malfunction were identified and rectified. However the present survey is to reduce unnecessary exposure to patients (who have been undergoing radio diagnostic procedure in various centers every day), radiation occupancy, and public, which also provide pipeline of Diagnostic Reference Levels (DRLs) and to create a new protocol for diagnostic instruments.

II. Images:

III. Reference:

- Delis H, Christaki K, Healy B, Loreti G, Poli GL, Toroi P and Meghzifene A, 2017. Moving beyond quality control in diagnostic radiology and the role of the clinically qualified medical physicist. Phys Med.41:104-108.

- Eija Metsälä, Anja Henner and Marja Ekholm, 2014. Quality assurance in digital dental imaging: a systematic review. Acta Odontologica Scandinavica.72: 362–371.

- Singh TD, Jayaraman T and Arunkumar Sharma B, 2017. Assessment of radiological protection systems among diagnostic radiology facilities in North East India. J Radiol Prot. 37(1):68-83.

IV. Biography:

C. Senthamil Selvan obtained his M.Sc. in Medical Physics, M.B.A. in Hospital Management and pursuing his Ph.D. in Medical Physics (dissertation submitted) from Bharathiar University. He has been serving as a peer reviewer of APES and IJRR journal. He has received numerous honors and awards such as ICTP visiting fellow, Italy in 2018, Swami Vivekananda Youth Achievement Award in 2018, Researcher and Educational Services Award in 2018, National Best Teacher Award, India in 2017, and Junior Research Fellow of DST-SERB research project in 2013. His research has been focused on medical imaging and radiation safety in diagnostic and therapeutic.

V. Presenting author details

Full Name: C. Senthamil Selvan

Contact number: +91-9047783116

Twitter account: https://twitter.com/PSSP_SERVICE

Linked In account: https://www.linkedin.com/in/c-senthamil-selvan-1a181474/

Session name/ number: -----------

Category: Oral Presentation

Passport Number: M6971280